# Submission ID #: 70318

# Scriptwriter Name: Pallavi Sharma

# Project Page Link: https://review.jove.com/files_upload.php?src=21447038

Title: Robotically Delivered fMRI-Guided Personalized Transcranial Magnetic Stimulation Therapy for Treatment-Resistant Depression

# Authors and Affiliations:

Bjorn Burgher1,2, Robin Cash3, Andrew Zalesky 3, Vinh Nguyen2, Luke Hearne1, Luca Cocchi1

1 QIMR Berghofer Medical Research Institute

2 Queensland Neurostimulation Centre

3 University of Melbourne

☐   All author names and affiliations are correct (city/state/country information not included in video title page).

# Corresponding Authors:

Bjorn Burgher		(bjorn.burgher@qimrb.edu.au)

# Email Addresses for All Authors:

Robin Cash		(robin.cash@unimelb.edu.au)

Andrew Zalesky 	(Andrew.zalesky@unimelb.edu.au)

Vinh Nguyen		(thaivinhnguyen@gmail.com)

Luke Hearne		(luke.hearne@qimrb.edu.au)

Luca Cocchi		(luca.cocchi@qimrb.edu.au)

Bjorn Burgher		(bjorn.burgher@qimrb.edu.au)

## Author Questionnaire

1. Microscopy: Does your protocol require the use of a dissecting or stereomicroscope for performing a complex dissection, microinjection technique, or something similar?  Enter Yes or No.

If Yes, can you record movies/images using your own microscope camera?

Enter Yes or No.

If your protocol involves microscopy but you are not able to record movies/images with your microscope camera, JoVE will need to use our scope kit.

If your microscope does not have a camera port, the scope kit will be attached to one of the eyepieces and you will have to perform the procedure using one eye.

Enter make and model of microscope.

If a dissection or stereo microscope is required for your protocol, please list all shots from the script that will be visualized using the microscope (shots are indicated with the 3-digit numbers, like 2.1.1, 2.1.2, etc.).

Click here to list microscope shots, using the shot numbers from the protocol section of the video script.

2. Software: Does the part of your protocol being filmed include step-by-step descriptions of software usage?  Enter Yes or No.

If Yes, we will need you to record using screen recording software.

We recommend using the screen capture program OBS. JoVE’s tutorial for using OBS Studio is provided at this link: https://review.jove.com/v/5848/screen-capture-instructions-for-authors?status=a7854k

As these files are necessary for finalizing your script, please upload all screen-captured video files to your project page as soon as possible.

3. Filming location: Will the filming need to take place in multiple locations?   Enter Yes or No.

If Yes, how far apart are the locations? Click to enter distance between locations.

4. Testimonials (optional): Would you be open to filming two short testimonial statements live during your JoVE shoot? These will not appear in your JoVE video but may be used in JoVE’s promotional materials. Enter Yes or No.

If Yes, please provide the full name and position (e.g., Director of [Institute Name], Senior Researcher [University Name], etc.) of the author willing to participate.

Enter full author name, Enter author position

To ensure that your script can be filmed in one day, the protocol sections are cumulatively restricted to 55 shots (shots are the 3-digit numbers like 2.1.1, 2.1.2…etc)

Current Protocol Length

Number of Steps:  33

Number of Shots:  58

# Introduction

Videographer: Obtain headshots for all authors available at the filming location.

Answers to these questions will become interview statements that you will deliver on camera.

- Introduction: Answer the 1st REQUIRED prompt and 1 additional question; Conclusion: Answer any 3 questions. No more than 5 interview statements will be included in the video.

- Enter the full name of the author who will deliver the statement.

- Speak naturally and avoid reading the lines.

- Limit the length of each statement to 20 words or fewer.

- Answers will be edited for length, clarity, and consistency with journal style guidelines.

Authors: Please note that the questions will not appear on screen. Please answer in stand-alone full sentence with sufficient context

INTRODUCTION:

REQUIRED: Summarize your research focus and the main questions it aims to answer.

  - Enter author name.: Click here to answer question. Please write in a style that you will be comfortable memorizing and speaking aloud. Limit length to 20 or fewer words.

Describe the research gap or limitation in existing methods and how this protocol addresses them.

  - Enter author name: Click here to answer question. Please write in a style that you will be comfortable memorizing and speaking aloud. Limit length to 20 or fewer words.

Describe the research contexts or systems in which this protocol can be applied.

  - Enter author name: Click here to answer question. Please write in a style that you will be comfortable memorizing and speaking aloud. Limit length to 20 or fewer words.

CONCLUSION:

What does this protocol allow researchers to study or measure?

  - Enter author name: Click here to answer question. Please write in a style that you will be comfortable memorizing and speaking aloud. Limit length to 20 or fewer words.

Describe the most important consideration or challenge when performing this protocol.

  - Enter author name: Click here if you choose this question. Please write in a style that you will be comfortable memorizing and speaking aloud. Limit length to 20 or fewer words.

Are there any additional methods or analyses that can be performed following this procedure?

  - Enter author name: Click here if you choose this question. Please write in a style that you will be comfortable memorizing and speaking aloud. Limit length to 20 or fewer words.

How can future studies build on this work? Describe the questions or research directions that can be investigated next.

  - Enter author name: Click here if you choose this question. Please write in a style that you will be comfortable memorizing and speaking aloud. Limit length to 20 or fewer words.

Videographer: Obtain headshots for all authors available at the filming location.

# Ethics Title Card

This research has been approved by the local Human Research Ethics Committee

# Protocol

Please review this section to make sure that it accurately describes your protocol. Use Track Changes when making edits or revisions.

- The two-digit steps (e.g., 2.1., 2.2.) with purple font are the narration.  JoVE is responsible for the narration of the protocol and results.

- Red italics are pronunciation guides indicating how the word will be spoken.

- Filming should take no more than 10 minutes per step. If a step takes more than 10 minutes, prepare the product for that step in advance.

- The three-digit shots (e.g., 2.1.1., 2.2.2.) are the actions that the videographer will capture.

Videographer: Please always record an audio slate while filming each shot mentioned in the shotlist.

- Screen patients for MRI safety exclusions (metal implants, claustrophobia, pregnancy)

- Assess baseline symptoms using MADRS, HAM-A, and HADS

- Acquire MRI scans (T1 + resting-state fMRI + field maps) using 3T scanner

Video editor: Please use a title card to show this information.

- Resting-state Functional Connectivity Analysis

Demonstrator: Click here to enter name of demonstrator(s)

If the same person is the demonstrator throughout, mention them once here and remove the "Demonstrator" field from the other sections; if the demonstrator changes, retain the field in the respective sections.

  - After converting the patient’s MRI (M-R-I) DICOM (Di-Com) files into NIFTI (Nifti) format and performing RSFC (R-F-S-C) analysis [1], identify the standard space brain coordinates located at the center of gravity of the left dorsolateral prefrontal cortex or DLPFC (D-L-P-F-C) anti-correlated cluster in x, y, z format [2-TXT]. Obtain the standard space DLPFC-SGC (D-L-P-F-C-S-G-C) functional connectivity map produced in NIFTI format [3].

    - WIDE: Talent in front of the computer with patients file showing RFSC analysis.

    - SCREEN: Display the standard space brain coordinates at the center of gravity of the left dorsolateral prefrontal cortex anti-correlated cluster or clusters in x, y, z format. TXT: SGC: Subgenual cingulate cortex

Authors: Please create screen capture videos of the shots labeled as SCREEN, create a screenshot summary, and upload the files to your project page as soon as possible: https://review.jove.com/files_upload.php?src=21447038

    - SCREEN: Show the standard space dorsolateral prefrontal cortex-subgenual cingulate functional connectivity map in Neuroimaging Informatics Technology Initiative format. TXT: NIFTI: Neuroimaging Informatics Technology Initiative

  - Identify the standard space brain coordinate corresponding to the right abductor pollicis brevis motor area to approximate the target for resting motor threshold or RMT (R-M-T) testing [1].

    - SCREEN: Display the standard space brain for the right abductor pollicis brevis motor area primary motor cortex.

  - Warp the DLPFC and M1 coordinates and DLPFC-SGC functional connectivity map from standard space to the patient’s native space using transforms generated by the initial preprocessing pipeline prepared earlier [1].

    - SCREEN: Show the DLPFC and M1 coordinates and DLPFC-SGC functional connectivity  map being warped from standard space  to the patient’s native space.

  - Provide the psychiatrist performing the TMS targeting with the native space brain coordinates for DLPFC and M1 and DLPFC-SGC functional connectivity map [1].

    - Talent shows the native space brain coordinates and connectivity map to the psychiatrist.

- Preparation for Neuronavigation

Demonstrator: Click here to enter name of demonstrator(s)

  - To navigate to the storage location of the patient’s native T1w (T-One-W) images, click File and select New Session in the neuronavigation [1]. Create a session folder and select either DICOM (.dcm) or nifti (.nii) (Di-Com-Dot-D-C-M or NIFTI-Dot-N-I-I), depending on the storage format of the T1w images [3].

    - SCREEN: Show the neuronavigation software interface and click File > New Session to navigate to the storage location of the patient’s native T1-weighted images.

    - SCREEN: Show creation of a new session folder within the selected storage location. Show selection between DICOM or NIfti.

  - Adjust the surface thresholds on the scalp and brain and predefine anatomical landmarks on the T1w image using the menu bar options in the neuronavigation software [1].

    - SCREEN: Show the neuronavigation software menu bar with options for adjusting surface thresholds and defining anatomical landmarks on the T1-weighted image.

  - Using the Adjust Scalp Threshold option, adjust the scalp surface threshold to inflate or deflate [1]. Ensure minimization of epidural space or scalp overinflation from the scalp surface definition [2].

    - SCREEN: Show adjustment of the scalp surface threshold using the Adjust Scalp Threshold option.

    - SCREEN: Show the scalp surface definition with minimized epidural space and no overinflation.

  - To segment brain tissues from non-brain tissues, using the brain segmentation option, select a large white matter tract and click Calculate so the software determines relative contrast values of brain and non-brain tissue [1]. Then, check that the defined brain surface adequately aligns with the grey matter surface [2].

    - SCREEN: Show the brain segmentation options, show selection of a large white matter tract and clicking Calculate.

    - SCREEN: Show the defined brain surface aligned with the grey matter surface.

  - Now, using the Patient Registration option, predefine anatomical landmarks [1]. Mark the nasion, right tragus, and left tragus onto the patient’s native-space T1w image in preparation for coregistration with the patient’s actual scalp landmarks [2].

    - SCREEN: Show selection of the Patient Registration option in the neuronavigation software.

    - SCREEN: Show marking of the nasion, right tragus, and left tragus on the patient’s native-space T1-weighted image.

  - Next, enter the native-space DLPFC target and the M1 coordinates in x, y, z format using right-anterior-superior orientation with the planning option in the menu bar [1].

    - SCREEN: Show entry of the dorsolateral prefrontal cortex and primary motor cortex coordinates in the planning option using right-anterior-superior orientation.

  - Reposition the DLPFC target to the nearest gyrus if it falls within a sulcus [1]. To verify that the repositioned brain coordinate remains within the selected anti-correlated cluster, overlay the native space DLPFC-SGC functional connectivity map onto the patient’s native space T1w image [2] and enter the coordinates of the repositioned brain target in the MRI visualization software [3].

    - SCREEN: Show repositioning of the dorsolateral prefrontal cortex target from a sulcus to the nearest gyrus.

    - SCREEN: Show overlay of the dorsolateral prefrontal cortex-subgenual cingulate functional connectivity map on the native-space T1-weighted image.

    - SCREEN: Entry of the repositioned coordinates.

  - Reposition the primary motor cortex coordinate to the precentral gyrus if necessary [1].

    - SCREEN: Show repositioning of the primary motor cortex coordinate to the precentral gyrus.

  - Calculate the perpendicular entry points of the TMS coil on the scalp for the DLPFC and M1 coordinates [1]. Ensure the approach angle of the TMS coil relative to the left side of the head is set to 45 degrees during calculation of the entry point [2].

    - SCREEN: Show calculation of perpendicular scalp entry points for the dorsolateral prefrontal cortex and primary motor cortex coordinates.

    - SCREEN: Show the transcranial magnetic stimulation coil angle set to 45 degrees relative to the left side of the head.

  - Overlay a 6 by 4 entry or target grid spaced at 5 by 5 millimeter intervals in the entry or target option in the control area to create additional targets for RMT testing [1].

    - SCREEN: Show overlay of a 6 by 4 entry and target grid with 5 by 5 millimeter spacing in the control area for resting motor threshold testing.

- Preparation of the Robotic Arm Workflow

Demonstrator: Click here to enter name of demonstrator(s)

  - Once the patient is seated in a motorized treatment chair with a neck rest, recline the patient and raise the legs to a comfortable position [1-TXT].

    - Talent reclining the chair and raising the patient’s legs to a comfortable position. TXT: Ensure the neck rest adequately supports the head and neck

  - Using a headband or double-sided tape, fix the patient reference tracker to the patient’s right forehead [1]. Position the patient’s head using the motorized chair so that it is within the visual field of the neuronavigation camera [2].

    - Talent attaching the patient reference tracker to the patient’s right forehead using a headband or double-sided tape.

    - Talent adjusting the motorized chair to position the patient’s head within the visual field of the neuronavigation camera.

  - Then, return to the patient registration option in the menu bar to coregister the patient’s anatomical landmarks and scalp surface to the predefined anatomical landmarks and scalp surface on the T1w image [1]. Touch the pointer to the patient’s actual scalp landmarks and acquire each landmark by pressing the Green button on the foot or hand switch [2].

    - SCREEN: Show selection of the Patient Registration option in the neuronavigation software.

    - Talent touching the pointer to the patient’s scalp landmarks and pressing the green button on the foot or hand switch to acquire each landmark.

  - Touch the pointer to the scalp and gently run it across the patient’s scalp to acquire 300 scalp landmarks while pressing the green button on the foot or hand switch [1].

    - Talent moving the pointer across the patient’s scalp while pressing the green button to acquire multiple scalp landmarks.

  - Check the spatial alignment accuracy between the actual and predefined anatomical landmarks and scalp surface at the end of the patient registration procedure, reported as root mean square deviation in millimeters [1-TXT].

    - SCREEN: Show the root mean square deviation value displayed in the neuronavigation software. TXT: Ensure RMS < 3 mm; Repeat steps if needed

  - Reposition the patient in the motorized chair so the patient reference tracker is at an optimal distance and angle within the visual field of the neuronavigation camera [1]. Ensure the visual field of the neuronavigation camera can also detect one of the cobot reference points [2]. Check the visual field by selecting Check Tracking System in the menu bar [3].

    - WIDE: Talent adjusting the motorized chair to reposition the patient for optimal tracking by the neuronavigation camera.

    - Shot showing the neuronavigation camera view including both the patient reference tracker and a cobot reference point.

    - SCREEN: Show selection of Check Tracking System in the neuronavigation software to verify the visual field.

  - Enable manual mode in the cobot option in the control area to manually move the coil-mounted robotic arm into the workspace [1]. Click Move to Park position in the cobot option panel to move the coil into the neutral position [2].

    - SCREEN: Show enabling of manual mode in the cobot option within the control area.

    - SCREEN: Show clicking Move to Park position and the coil moving to the neutral position.

  - Manually position the cobot or adjust the motorized chair so the coil is directly above the patient’s head with a separation of several centimeters [1].

    - Talent adjusting the cobot or motorized chair to position the coil above the patient’s head with a small gap.

  - Now, calibrate the coil force sensor function in the cobot option within the control area of the neuronavigation software to trigger a sensor pop-up window with four pressure lights [1]. Place a finger on the coil and gradually increase and then decrease pressure so all four pressure lights turn on and then off in sequence [2].

    - SCREEN: Show activation of the coil force sensor calibration and display of the pop-up window with four pressure lights.

    - Talent pressing on the coil with a finger to sequentially activate and deactivate all four pressure lights.

- Resting Motor Threshold Testing

Demonstrator: Click here to enter name of demonstrator(s)

  - Place three electromyograph electrodes connected to the MEP monitor along the ventral length of the right abductor pollicis brevis [1]. Select the MEP protocol on the TMS machine and set the sensitivity to 200 microvolt per division [2]. Select Recall and then select Timing to display the MEP interface [3].

    - WIDE: Talent placing three electromyograph electrodes along the ventral length of the right abductor pollicis brevis.

    - SCREEN: Show selection of the motor evoked potential protocol and adjustment of sensitivity to 200 microvolt per division on the transcranial magnetic stimulation machine.

    - SCREEN: Show selection of Recall and then Timing to display the motor evoked potential interface.

  - To instruct the coil-mounted robotic arm to move to the first M1 target in the grid, select the target in the entry and target option panel in the control area [1]. Select Coil Alignment in the cobot option panel to position the coil on the patient’s scalp [2].

    - SCREEN: Show selection of the first primary motor cortex target in the entry and target option panel.

    - SCREEN: Show clicking Coil Alignment and positioning of the coil on the patient’s scalp.

  - Select the stimulator option in the control area of the neuronavigation software and set the initial stimulation intensity to 40 percent in the amplitude option [1]. Click Single Pulse to deliver a pulse while moving sequentially to each of the targets in the primary motor cortex grid [2-TXT].

    - SCREEN: Show selection of the stimulator option and setting the amplitude to 40 percent.

    - SCREEN: Show clicking Single Pulse while iterating through targets in the primary motor cortex grid.  TXT: Keep the patient’s hand and arm relaxed during RMT testing

  - Increase the stimulation intensity as needed to observe a MEP or visible thumb flexion [1]. Identify the target that produces the strongest MEP along with the largest amplitude flexion of the right thumb, fingers, or wrist [2].

    - SCREEN: Show increasing stimulation intensity and observing motor evoked potential responses on the electromyography display.

    - SCREEN: Split view showing motor evoked potential signals on the screen and visible flexion of the patient’s right thumb or fingers indicating the optimal target.

  - Select the Optimal M1 target in the entry and target panel and click Coil Alignment in the cobot option panel to position the coil on the patient’s scalp [1]. Adjust the stimulation intensity of single pulses at this target at set intervals until the RMP is determined [2]. Identify the RMP as the minimum stimulation intensity required to evoke 5 out of 10 movements with a motor evoked potential greater than 50 millivolt on electromyography [3].

    - SCREEN: Show selection of the optimal primary motor cortex target and clicking Coil Alignment to position the coil.

    - SCREEN: Show adjustment of stimulation intensity while delivering single pulses at the selected target.

    - SCREEN: Show identification of the resting motor threshold based on motor evoked potential responses on electromyography.

- Delivering Stimulation to the Personalized DLPFC Target

Demonstrator: Click here to enter name of demonstrator(s)

  - Instruct the patient to close the eyes, relax, and avoid focusing on any specific thought to maintain consistent brain states across fMRI and TMS [1].

    - Talent instructing the patient to close the eyes and relax in the treatment chair.

  - Instruct the coil-mounted robotic arm to move to the DLPFC target by selecting the target in the entry and target option panel in the control area [1]. Select Coil Alignment in the cobot option panel to position the coil on the patient’s scalp [2].

    - SCREEN: Show selection of the dorsolateral prefrontal cortex target in the entry and target option panel.

    - SCREEN: Show clicking Coil Alignment and positioning of the coil on the patient’s scalp.

  - Select the Programmed 3 Minute Intermittent Theta Burst Stimulation protocol in the menu option on the TMS machine [1]. Select Recall and then Timing to prepare the machine to deliver the 3 minute intermittent theta burst stimulation protocol [2].

    - SCREEN: Show selection of the programmed 3 minute intermittent theta burst stimulation protocol in the menu.

    - SCREEN: Show selection of Recall and Timing to prepare the stimulation protocol.

  - Select the Stimulator option in the control area and set the stimulation intensity to 80 percent of RMP for the first session, 100 percent for the second session, and 120 percent for subsequent sessions [1]. Click Start/Stop Trains (Start-Or-Stop-Trains) to deliver the full train of repetitive transcranial magnetic stimulation pulses [2].

    - SCREEN: Show setting of stimulation intensity relative to resting motor threshold in the stimulator option.

    - SCREEN: Show clicking Start/Stop Trains to deliver repetitive stimulation pulses.

  - Perform left dorsolateral prefrontal cortex stimulation daily for up to 20 to 30 weekday sessions [1].

    - Talent administering repeated daily transcranial magnetic stimulation sessions to the patient.

# Results

Please review this section to make sure that it accurately reflects your findings.

- This section will not be recorded by the videographer. It only includes the figures/tables from your manuscript (called LAB MEDIA).

- Use Track Changes when making edits or revisions. Ensure the voiceover length is below 200 words. Current word count: 85.

- Please note that the video cannot include voiceover without an accompanying visual.

- Results

  - The study demonstrated an overall response rate of 52% [1] and remission rate of 33% [2].

    - LAB MEDIA: Figure 1A. Video editor: Highlight the bar labelled “All” with the number 52 in the response section

    - LAB MEDIA: Figure 1A. Video editor: Highlight the bar labelled “All” with the number 33 in the remission section

  - The a priori subgroups showed distinct response [1] and remission rate profiles, indicating lower treatment responsiveness in the subgroups associated with comorbidity and bipolar depression [2].

    - LAB MEDIA: Figure 1A,B. Video editor: Show the subgroup response bars for G1, G2, and G3 so the different heights can be compared

    - LAB MEDIA: Figure 1A,B. Video editor: Show the subgroup remission bars for G1, G2, and G3 so the different heights can be compared

  - Pre-treatment, the group average functional connectivity of the SGC-DLPFC was minus 0.25 [1].

    - LAB MEDIA: Figure 2B. Video editor: Highlight the “Pre” point in the left graph at about -0.25

  - In the subset of patients who underwent pre- and post-TMS fMRI, a specific reduction in anti-correlated RSFC was observed from pre-treatment to post-treatment [1], with the negative resting-state functional connectivity approaching 0 [2].

    - LAB MEDIA: Figure 2B. Video editor: Highlight the left graph labelled “DLPFC - SGC” and show the line moving upward from “Pre” to “Post”

    - LAB MEDIA: Figure 2B. Video editor: Emphasize that the “Post” point in the left graph is closer to the dashed 0 line than the “Pre” poin