Screen Recordings
· 70644_ 11-25-59.mp4
· 3.5.4.1 (Click on Lab Chart and wait for it to load) 00:02-00:07
· 3.5.4.2 (There will be a green check mark under power labs and a red ‘X’ under lab chart compatible devices. Click OK.) 00:10-00:17
· 3.5.4.3 (Click ‘New’ on bottom left corner) 00:19-00:21
· 3.5.4.4 (Click ‘Set up at the top left, scroll to ‘devices and channels’ and select it. Change number of channels to ‘4’ and click OK. 00:22-00:32
· 3.5.4.5 (Right click arrow next to channel 1. Click ‘rename channel’. Rename channel to ‘1: PA Cuff Pressure’. Click OK.) 00:35-00:40
· 3.5.4.6 (Repeat renaming channels 2,3, and 4 with 2: RV Pressure, 3: RVSP, and 4: Heart Rate, respectively.) 00:42-00:56
· 3.5.4.7 (Right click empty space under channel 3 and select ‘cyclic measurements’. Click source and change to ‘Channel 2: RV Pressure’. Change measurement to ‘maximum’. This channel will now calculate right ventricle systolic pressure. Click OK.) 00:58-01:07
· 3.5.4.8 (Repeat process for Channel 4 Heart Rate but change measurement to rate) 01:08-01:18
· 3.5.4.9 (Repeat process for channel 1 and 2 but ensure no calculation is selected) 01:19-01:23
· 70644 11-29-43.mp4
· 4.1.1.1 (Click ‘Start’ on bottom right corner. Right click empty space under Channel 1 and select ‘Bridge Amp…’) 00:04-00:08
· 4.1.1.2 (A small screen will appear and the red line should be around 0) 00:08-00:10
· 4.1.1.3 (Turn tab on pressure transduce so that it is off to patient) This will need to be filmed.
· 4.1.1.4 (Connect sphygmomanometer to pressure transducer and inflate until device reads 80mmHg) This will need to be filmed.
· 4.1.1.5 (Click on units. The unit conversion screen will display. Click the ‘+’ button to zoom in until you see a line with a large peak.) 00:29-00:32
· 4.1.1.6 (Highlight the first flat portion then click the arrow next to ‘point 1’. Then enter zero in the second box next to the second arrow. Next, highlight the tall peak and click the arrow next to ‘point 2’. Put 80 next to the second box next to the second arrow. Change units to mmHg and click OK twice.) 00:32-00:53
· 4.1.1.7 (Repeat Process for channel 2 and then click start to begin recording values.) 02:42-2:44


