# Submission ID #: 71124

# Scriptwriter Name: Poornima G

# Project Page Link: https://review.jove.com/account/file-uploader?src=21692868

# Title: A Web Platform for Automated Image-Based Quantification of Human and Rodent Isolated Pancreatic Islets

# Authors and Affiliations:

# Sarah Suergiu1,2, Ivan Leontovyc1, Frantisek Saudek1, David Habart1

# 1Laboratory for Pancreatic Islets, Experimental Medicine Center, Institute for Clinical and Experimental Medicine

# 2Faculty of Science, Charles University

# Corresponding Authors:

# Sarah Suergiu		sues@ikem.cz

# Email Addresses for All Authors:

# Ivan Leontovyc	ivlo@ikem.cz

# Frantisek Saudek 	frsa@ikem.cz

# David Habart		daht@ikem.cz

# Sarah Suergiu		sues@ikem.cz

## Author Questionnaire

1. Microscopy: Does your protocol require the use of a dissecting or stereomicroscope for performing a complex dissection, microinjection technique, or something similar?  Yes

If Yes, can you record movies/images using your own microscope camera?

Yes

If your protocol involves microscopy but you are not able to record movies/images with your microscope camera, JoVE will need to use our scope kit.

If your microscope does not have a camera port, the scope kit will be attached to one of the eyepieces and you will have to perform the procedure using one eye.

Olympus CKX41

If a dissection or stereo microscope is required for your protocol, please list all shots from the script that will be visualized using the microscope (shots are indicated with the 3-digit numbers, like 2.1.1, 2.1.2, etc.).

Click here to list microscope shots, using the shot numbers from the protocol section of the video script.

2. Software: Does the part of your protocol being filmed include step-by-step descriptions of software usage?  Yes

If Yes, we will need you to record using screen recording software.

We recommend using the screen capture program OBS. JoVE’s tutorial for using OBS Studio is provided at this link: https://review.jove.com/v/5848/screen-capture-instructions-for-authors?status=a7854k

Please upload all screen-captured video files to your project page as soon as possible.

Videographer: Please record the computer screen for the shots labeled as SCREEN. The screen should be in focus and the fonts should be clearly visible.

3. Filming location: Will the filming need to take place in multiple locations?   No

4. Testimonials (optional): Would you be open to filming two short testimonial statements live during your JoVE shoot? These will not appear in your JoVE video but may be used in JoVE’s promotional materials. Yes

Current Protocol Length

Number of Steps:  33

Number of Shots:  76 (55 SC)

Interviews

Videographer: Obtain headshots for all authors available at the filming location.

Authors: Please note that the questions will not appear on screen. Please answer in stand-alone full sentence with sufficient context

We have chosen the 5 best answers

INTRODUCTION:

REQUIRED: Summarize your research focus and the main questions it aims to answer.

  - David Habart: We are developing accessible and accurate image analysis methods for islet quantification as an indirect measure of β-cell mass in transplantation.

    - INTERVIEW: Named talent says the statement above in an interview-style shot, looking slightly off-camera

Describe the research gap or limitation in existing methods and how this protocol addresses them.

  - David Habart: Existing methods rely on subjective thresholding that struggle with the inherent sample variability. We provide efficient deep learning-based image analysis with improved accessibility through an online service.

    - INTERVIEW: Named talent says the statement above in an interview-style shot, looking slightly off-camera

CONCLUSION:

What does this protocol allow researchers to study or measure?

  - David Habart: IsletNet analyses single 2D microscopic projections and returns islet count, size histogram, and volume estimates using several existing methods with side-by-side comparison.

    - INTERVIEW: Named talent says the statement above in an interview-style shot, looking slightly off-camera

Describe the most important consideration or challenge when performing this protocol.

  - David Habart: Reliable performance requires microscopy images consistent with those used during IsletNet training. If imaging conditions change, retraining with additional images is recommended.

    - INTERVIEW: Named talent says the statement above in an interview-style shot, looking slightly off-camera

How can future studies build on this work? Describe the questions or research directions that can be investigated next.

  - David Habart: IsletNet can support future standardized multicenter studies using isolated islets and unbiased comparisons with re-aggregated pseudo-islets and stem cell-derived islets.

    - INTERVIEW: Named talent says the statement above in an interview-style shot, looking slightly off-camera

Videographer: Obtain headshots for all authors available at the filming location.

# Ethics Title Card

This research has been approved by the Ministry of Health of the Czech Republic

Protocol

Videographer: Please always record an audio slate while filming each shot mentioned in the shotlist.

Authors: Our shot limit per video is 55. We are considering your case an exception after a discussion and allowing 20 extra shots. Please note that it will be difficult for us to accommodate more.

# Authors: Please have all reagents, samples, and instruments ready before the shoot so the shoot can proceed without delay, as this is a long script. If an end product requires a long incubation period, please prepare that sample in advance and have it ready for the shoot.

- Software Setup

  - Prepare the software environment for image acquisition, storage, and analysis, and launch the camera software [1]. Create folders for the expected image files [2]. Open the IsletNet website and confirm the displayed training date [3].

    - WIDE: Talent opening the live-view window of the Quick Photo software on the computer.

    - SCREEN: show the folder with subfolders labelled High purity, Medium purity, Low purity, and Calibrator.

    - SCREEN: opening the IsletNet Home page and highlighting the training date.

Videographer: Please record the computer screen for the shots labeled as SCREEN. The screen should be in focus and the fonts should be clearly visible.

  - Access the image submission page through the Clinical Isolation link or the Upload button [1]. Start the session by entering the operator initials, and current session title, [2].

    - SCREEN: clicking Clinical Isolation and opening the batch submission page.

    - SCREEN: entering the operator initials, current session.

Authors: Our videographer will film the computer screen for the shots labeled "SCREEN". But I also strongly suggest that you keep the screen recording software running on the computer throughout the procedure so we can fall back on it if we have issues.

- Islet Sampling and Dithizone Staining

  - After islet purification, place the 250-milliliter conical tubes containing the high-, medium-, and low-purity fractions into a cooling box inside the laminar flow hood [1].

    - Talent placing the cooling box, labelled tubes, pipettes, dithizone working solution, and four marked Petri dishes inside the hood.

  - Dispense 25 microliters of 2.5 milligrams per milliliter dithizone working solution into the center of four marked Petri dishes designated for high-purity replicate samples, avoiding bubble formation [1-TXT]. Replace the tip after each dish [2].

    - Talent dispensing dithizone into the center of the dishes. TXT: Label the dishes as H1, H2, H3 and H4; Prepare the dithizone in an alkaline alcoholic solution

    - Talent replacing the pipette tip.

  - Immediately before sampling, gently invert the high-purity tube five times [1] and open the lid [2].

    - Talent inverting the high-purity tube.

    - Talent opening the lid.

  - Using a large-bore pipette tip, collect one replicate sample from the midpoint of the suspension [1]. Dispense the islet suspension into the dithizone drop and mix by pipetting up and down three times, using only the first stop of the plunger to avoid bubble formation [2]. Replace the tip after dispensing [3].

    - Talent collecting the islet suspension using a large-bore pipette tip.

    - Shot of dispensing the suspension into the dithizone drop and mixing slowly three times.

    - Talent discarding the pipette tip and attaching another tip.

  - Before collecting each additional replicate sample, gently invert the original tube five times again to maintain representative sampling [1]. Repeat the sampling and dispensing procedure for the next dish [2].

    - Talent inverting the tube before the next replicate.

    - Talent dispensing another sample into the dithizone drop in the next dish.

  - Transfer the tray containing the four high-purity replicate samples to the microscope bench [1].

    - Talent transferring the sample tray to the microscope bench with the computer, camera software, and IsletNet visible.

  - After 1- to 2-minutes of incubation, place dish H1 under the microscope [1] and visually monitor the staining until the islets appear bright red [2].

    - Talent placing dish H1 on the microscope stage.

    - Close-up or SCREEN shot of well-stained islets.

  - Then, add 4 milliliters of albumin-supplemented PBS directly into the drop in the dish H1 while avoiding bubble formation [1]. Swirl each dish vigorously to disperse the dithizone and keep the tissue near the center [2].

    - Talent adding albumin-supplemented PBS to the dish.

    - Talent vigorously swirling the dish.

  - Immediately, inspect dish H1 under the microscope and rotate the dish to locate any islets outside the camera field [1]. Move them toward the center using fluid generated by a 200-microliter pipette tip, without direct contact with the islets [2]. Place the dish on the bench without disturbing the islets [3-TXT].

    - Talent inspecting and rotating the dish under the microscope.

    - SCREEN: moving displaced islets toward the center using the pipette tip.

    - Talent placing the dish back on the bench. TXT: Repeat the procedure for dishes H2-H4 without any delay

Authors: We will film and show only 1 replicate as per our journal guidelines. I have added the above text in bold. This text will be displayed on screen so that the viewers know that they have to repeat the procedure for other replicates

- Imaging the Stained Islets

  - Immediately after staining, place dish H1 under the microscope [1]. Turn off the room lights [2] and adjust the halogen light to the standardized intensity [3].

    - Talent placing the H1 dish under microscope.

    - Talent turning off the room lights.

    - Talent adjusting the halogen light source.

  - REMOVE, was not done properly on medium purity fraction sample.

  - Select the pre-saved automatic camera settings and acquire the image [1]. Save it in the high-purity folder. Use JPG format when preferred to speed up analysis, and include magnification, dilution and pellet information in the file name [2].

    - SCREEN: selecting the saved camera settings and clicking Acquire.

    - SCREEN: saving the image in the high-purity folder with the magnification and other details included in the file name.

  - Repeat the sampling, staining, and imaging process for the medium- and low-purity fractions [1]. Save the images in their respective folders [2]. Note the increasing exocrine tissue content in lower-purity fractions [3].

    - Talent placing the medium-purity sample dish on the microscope stage.

    - SCREEN: Saving the medium-purity images in the correct folder.

    - SCREEN: Cursor hovering over the medium-purity image and highlighting exocrine tissue.

  - Next, turn on transmitted light, align the micrometer scale with the long edge of the camera frame, focus the image, and acquire an image of the stage micrometer under the same magnification [1]. Save it in the calibrator folder, including the magnification information in the file name [2].

    - SCREEN: Aligning the micrometer scale, focusing and acquiring the stage micrometer image.

    - SCREEN: Saving the calibration image in the Calibrator folder.

- Clinical Isolation Module: On-Site Use

  - Upload the high-purity sample images using the Add Images button or by dragging the files from the folder [1].

    - SCREEN: selecting the high-purity image files for upload.

  - Adjust the calibration according to the magnification specified in the uploaded image file names [1]. Open the pixel size calibration link and upload the corresponding stage micrometer image [2].

    - SCREEN:: ACTIONS SWITCHED uploading or dragging the stage micrometer image into the calibration interface.

    - SCREEN:  ACTIONS SWITCHED Pointing at the sample's image file names and hovering the cursor over the magnification in the name.

  - Position the endpoints of the red measurement line across a substantial distance on the digital ruler [1]. Enter the selected distance in micrometers [2] and click Submit, accepting the calculated pixel size [3].

    - SCREEN: drawing the measurement line across the ruler.

    - SCREEN: entering the known distance in micrometers.

    - SCREEN: accepting the calculated pixel size.

  - Back in the batch submission page, enter the minimum islet size threshold, conventionally 50 micrometers [1].

    - SCREEN: entering the minimum islet size threshold.

  - Enter the sample dilution from the image file name [1], the pellet volume estimated from the original conical tube [2], and batch ID.

    - SCREEN: entering the sample dilution.

    - SCREEN: entering the pellet volume in milliliters.

    - SCREEN: enter batch ID.

  - Click Submit Batch to launch the image analysis [1]. Wait till both the numerical and graphical results appear in the Inspection window [2].

    - SCREEN: clicking Submit Batch.

    - SCREEN: the numerical and graphical results appearing.

  - Review the preliminary individual sample values and batch averages for islet count, volume, and purity. For more details on the current image summary, review the red rectangle on the left [1].

    - SCREEN: the cursor highlighting the file names, individual values, and calculated averages. Then hover the cursor over the red rectangle on the left.

- Visual Quality Control

  - Open the Visual Quality Control window by clicking the image [1]. Enable the legend to view keyboard shortcuts, including zoom and view-toggle controls [2]. Zoom in and out using the  “i” and “o” (eye and oh) keys [3].

    - SCREEN: opening the Visual Quality Control window.

    - SCREEN: enabling the legend and highlighting the zoom indicator.

    - SCREEN: zooming in using the keyboard.

  - Evaluate the automatic contours by toggling between the original image view and contour view, preferentially using the number keys to remain focused in the islet border [1]. Toggle to the islet segmentation view to confirm that islets below the counting threshold appear in the segmentation view without corresponding contours [2]. Use the down-arrow key to proceed to the next sample image [3].

    - SCREEN: toggling between the original image, contour view.

    - SCREEN: toggling between contour and segmentation views.

    - SCREEN: pressing the down-arrow key to move to the next image

  - MOVE between 6.6. and 6.7. Exists as screen recording only. After validating the high-purity batch, click on Add Another Batch [1]. Upload the medium-purity fraction images [2]. Enter the dilution and pellet volume [3]. Adjust calibration if necessary [4]. Type in the batch ID and submit it for analysis and wait till the Inspection window opens [5]. Validate no exocrine tissue was mistaken for islets [6].

    - SCREEN: clicking Add Another Batch.

    - SCREEN: uploading medium-purity images.

    - SCREEN: entering dilution and pellet volume.

    - SCREEN: highlighting the pixel size field.

    - SCREEN: submitting the batch.

    - SCREEN: toggling between contour and original views.

  - If adjacent islets are not correctly separated [1], left-click outside the islets to initiate a separation line [2]. Add intermediate left-clicks to guide the line between adjacent islets crossing the contour [3]. Complete the line with a right-click [4]. To remove incorrect manually added separation lines, right-click any point on the line and select the remove option from the pop-up menu [5]

    - SCREEN:

    - SCREEN: initiating a separation line outside the islets by right clicking outside islets.

    - SCREEN: adding intermediate clicks between adjacent islets.

    - SCREEN: completing the separation line outside the islets by right clicking.

    - SCREEN: right-clicking a separation line and selecting the remove option.

  - Submit the separation line and wait for the Inspection window to show the updated analysis [1], marked with the blue check [2].

    - SCREEN: submitting the separation line and updating the result.

    - SCREEN: Cursor hovering the blue check.

  - Proceed sequentially with the next sample images in the Visual Quality Control window. Use the switch at the top to exclude any invalid image containing undetected, partially detected, or spurious islets [1]. Also, exclude images in which exocrine tissue, background artifacts, or bubbles are incorrectly classified as islets [2].

    - SCREEN: show the next image with incorrect contouring or misclassified objects.

    - SCREEN: excluding the invalid image from analysis.

  - Use Edit Batch to correct batch information if needed [1]. Modify available parameters, do not change the session code, and do not reuse the names of removed images [2].

    - SCREEN: clicking Edit Batch to reopen the configuration window.

    - SCREEN: Modifying the parameters.

- Report Download

  - After reviewing and correcting all the batches, proceed to the session summary [1]. View the sums of islet count, volume, purity, and pellet values, together with the corresponding batch averages calculated from the included images [2].

    - SCREEN: clicking Summary.

    - SCREEN: the session summary with total and mean batch values.

  - Download the Complete PDF report for review and records [1].

    - SCREEN: downloading the report.

  - Finally, open the report to validate it. Confirm whether manual editing was used, and verify the batch results, including the number of images used and the pixel sizes [1]. Review the yield, pellet, islet size index, and islet size histogram mainly [2]. If the Report is valid, close the current Islet-Net session. The data cannot be retrieved [3].

    - SCREEN: Opening the report and showing the parameters.

    - SCREEN: Hover the cursor over the yield, pellet, islet size index, and islet size histogram.

    - SCREEN: Closing IsletNet

Authors: Skipping the typical results part that was added initially in the script as per our discussion.

Also, we are keeping the high-quality prerecorded video you submitted as a potential backup.

We can discuss and choose the footage that best represents the protocol and is of the highest quality.

Please note: All videos should fit into the narration as outlined in this version of the script.

NOTE to the Authors:

You have two options for recording and submitting the SCREEN footage:

Option 1: Uploading Separate Clips (Highly Recommended)

Please upload a separate video file for each SCREEN shot (tagged as SCREEN), naming the file according to the shot number (e.g., 2.4.3, 1.5.1, 3.2.1, 3.1.3, etc.).

- Each clip must be a maximum of 20 to 25 seconds to match the voice narration.

- If an action takes longer, provide only the best, most representative 20-25 second segment.

- Please do not provide longer clips as we won’t be able to use them.

Option 2: Uploading a Single Continuous Video

If you prefer to record all steps together and upload a single video file, that is acceptable. In this case, you must write the precise timestamp for each shot within the script.

For example, if the action described in shot 2.13.1 is present in the file xxx.mp4 and occurs between timestamps 00:30 and 00:45 , then write the following in the script after each shot description or as a comment bubble.

(write the file name and timestamp after each shot): [File name] [Start Time]-[End Time]

- Example 	2.13.1  SCREEN: performing ...........action. xxx.mp4 00:30-00:45

2.13.2, SCREEN: cutting the…..........action. xxx.mp4 01:10-01:20

- The duration between the start and end times for each shot must be limited to 20 to 25 seconds.

- If a shot is longer than 25 seconds, choose two short segments (one at the beginning and one at the end) that collectively do not exceed 25 seconds.

- Example: 3.1.1 SCREEN: inserting a needle... xxx.mp4 00:00-00:10 and 02:30-02:40