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Author Questionnaire
1. Microscopy: Does your protocol require the use of a dissecting or stereomicroscope for performing a complex dissection, microinjection technique, or something similar?  Enter Yes or No.  
If Yes, can you record movies/images using your own microscope camera?
Enter Yes or No.  
If your protocol involves microscopy but you are not able to record movies/images with your microscope camera, JoVE will need to use our scope kit. 
If your microscope does not have a camera port, the scope kit will be attached to one of the eyepieces and you will have to perform the procedure using one eye.
Enter make and model of microscope.
If a dissection or stereo microscope is required for your protocol, please list all shots from the script that will be visualized using the microscope (shots are indicated with the 3-digit numbers, like 2.1.1, 2.1.2, etc.).
Click here to list microscope shots, using the shot numbers from the protocol section of the video script.

2. Software: Does the part of your protocol being filmed include step-by-step descriptions of software usage?  Enter Yes or No.
If Yes, we will need you to record using screen recording software.
We recommend using the screen capture program OBS. JoVE’s tutorial for using OBS Studio is provided at this link: https://review.jove.com/v/5848/screen-capture-instructions-for-authors?status=a7854k
As these files are necessary for finalizing your script, please upload all screen-captured video files to your project page as soon as possible.

3. Filming location: Will the filming need to take place in multiple locations?   Enter Yes or No.
If Yes, how far apart are the locations? Click to enter distance between locations.

4. Testimonials (optional): Would you be open to filming two short testimonial statements live during your JoVE shoot? These will not appear in your JoVE video but may be used in JoVE’s promotional materials. Enter Yes or No.

If Yes, please provide the full name and position (e.g., Director of [Institute Name], Senior Researcher [University Name], etc.) of the author willing to participate.
Enter full author name, Enter author position

To ensure that your script can be filmed in one day, the protocol sections are cumulatively restricted to 55 shots (shots are the 3-digit numbers like 2.1.1, 2.1.2…etc)

Current Protocol Length
Number of Steps:  26
Number of Shots:  50

Interviews
Videographer: Obtain headshots for all authors available at the filming location. 

Answers to these questions will become interview statements that you will deliver on camera.
· Introduction: Answer the 1st REQUIRED prompt and 1 additional question; Conclusion: Answer any 3 questions. No more than 5 interview statements will be included in the video.
· Enter the full name of the author who will deliver the statement.
· Speak naturally and avoid reading the lines.
· Limit the length of each statement to 20 words or fewer.
· Answers will be edited for length, clarity, and consistency with journal style guidelines.

Authors: Please note that the questions will not appear on screen. Please answer in stand-alone full sentence with sufficient context

INTRODUCTION:

REQUIRED: Summarize your research focus and the main questions it aims to answer.
1.1. Enter author name.: Click here to answer question. Please write in a style that you will be comfortable memorizing and speaking aloud. Limit length to 20 or fewer words.

Describe the research gap or limitation in existing methods and how this protocol addresses them.
1.2. Enter author name: Click here to answer question. Please write in a style that you will be comfortable memorizing and speaking aloud. Limit length to 20 or fewer words.
Describe the research contexts or systems in which this protocol can be applied.
1.3. Enter author name: Click here to answer question. Please write in a style that you will be comfortable memorizing and speaking aloud. Limit length to 20 or fewer words.
CONCLUSION:
What does this protocol allow researchers to study or measure?
1.4. Enter author name: Click here to answer question. Please write in a style that you will be comfortable memorizing and speaking aloud. Limit length to 20 or fewer words.

Describe the most important consideration or challenge when performing this protocol.
1.5. Enter author name: Click here if you choose this question. Please write in a style that you will be comfortable memorizing and speaking aloud. Limit length to 20 or fewer words.

Are there any additional methods or analyses that can be performed following this procedure?
1.6. Enter author name: Click here if you choose this question. Please write in a style that you will be comfortable memorizing and speaking aloud. Limit length to 20 or fewer words.

How can future studies build on this work? Describe the questions or research directions that can be investigated next.
1.7. Enter author name: Click here if you choose this question. Please write in a style that you will be comfortable memorizing and speaking aloud. Limit length to 20 or fewer words.


Videographer: Obtain headshots for all authors available at the filming location.

Ethics Title Card
This research has been approved by the Institutional Animal Care and Use Committee (IACUC) at Massachusetts General Hospital 


Protocol  
[bookmark: _Hlk188263998]Please review this section to make sure that it accurately describes your protocol. Use Track Changes when making edits or revisions.
· The two-digit steps (e.g., 2.1., 2.2.) with purple font are the narration.  JoVE is responsible for the narration of the protocol and results.
· Red italics are pronunciation guides indicating how the word will be spoken. 
· Filming should take no more than 10 minutes per step. If a step takes more than 10 minutes, prepare the product for that step in advance.
· The three-digit shots (e.g., 2.1.1., 2.2.2.) are the actions that the videographer will capture. 
Videographer: Please always record an audio slate while filming each shot mentioned in the shotlist.
2. Setting up the Resonance Raman Spectroscopy (RRS) System 
Demonstrator: Click here to enter name of demonstrator(s) 

If the same person is the demonstrator throughout, mention them once here and remove the "Demonstrator" field from the other sections; if the demonstrator changes, retain the field in the respective sections.


2.1. To begin, connect the spectrometer unit to the direct current power supply [1] and connect it to the universal serial bus port on the computer running the analysis software using the supplied cables [2]. Identify the spectrometer unit with the round plug that contains alignment indentations [3].
2.1.1. WIDE: Talent connecting the spectrometer unit to the direct current power supply. Videographer: Please obtain CU shot as well for this shot.
2.1.2. Talent connecting the spectrometer unit to the universal serial bus port on the computer using the supplied cable.
2.1.3. Close-up of the spectrometer unit showing the round plug with alignment indentations.
2.2. Align the indentations of the round plug with those of the cable connected to the power supply and universal serial bus port [1]. Connect the ports without cleaning them beforehand [2].
2.2.1. Talent aligning the indentations of the round plug with the cable connector.
2.2.2. Talent connecting the ports together.
2.3. Place the laser probe in the probe holder [1]. Then adjust the probe's position at the start of the experiment [2].
2.3.1. Talent placing the laser probe in the probe holder.
2.3.2. Talent adjusting the position of the laser probe.
2.4. Open the spectrometer control software [1]. Observe the graph showing percent Resonance Raman reduced mitochondrial ratio or 3RMR (Three-R-M-R), tissue oxygen saturation, and myoglobin saturation upon opening the spectrometer control software [2]. Click the Settings tab at the bottom left of the page [3].
2.4.1. SCREEN: Spectrometer control software being opened.
Authors: Please create screen capture videos of the shots labeled as SCREEN, create a screenshot summary, and upload the files to your project page as soon as possible:  https://review.jove.com/files_upload.php?src=21743498
2.4.2. SCREEN: Graph displaying the percent 3RMR, tissue oxygen saturation, and myoglobin saturation metrics in the spectrometer control software.
2.4.3. SCREEN: Settings tab being selected at the bottom left of the page.
2.5. Under the Info tab, enter the session information in Session Name [1] and any necessary details in Session Description [2].
2.5.1. SCREEN: Info tab being opened. Session information being entered into Session Name.
2.5.2. SCREEN: Details being entered into Session Description.
2.6. Select Continuous Mode as the data collection mode for the experiment. Confirm that spectra will be recorded continuously until the stop button is pressed [2].
2.6.1. SCREEN: Continuous Mode being selected as the data collection mode. Continuous recording mode being displayed in the software.
2.7. Once all session information is entered, click the System tab [1].
2.7.1. SCREEN: System tab being selected.
2.8. Under Settings File, select the protocol appropriate for the project [1].
2.8.1. SCREEN: Settings File menu being opened and a protocol being selected.
2.9. To create a new settings file, adjust the required settings parameters. Then, click the Save button to save the parameters as a separate settings file [1].
2.9.1. SCREEN: Settings parameters being modified. Save button being selected and the settings file being saved.
2.10. To import a new settings file, click the drop-down menu next to Default (Mod) (Default -Mod) [1]. Select Import New Settings File and click the new settings file to import it into the system [2].
2.10.1. SCREEN: Drop-down menu next to Default (Mod) being opened.
2.10.2. SCREEN: Import New Settings File being selected. New settings file being selected and imported into the system.
2.11. In the Regression Libraries section, select and confirm the libraries for data acquisition and analysis. The libraries for reduced and oxidized mitochondria, complex III, complex IV, hemoglobin, and myoglobin were recorded before the experiment [1].
2.11.1. SCREEN: Regression Libraries section being opened. The library combination is being selected for data acquisition and analysis.
2.12. Select and confirm the preferred Window Width Mode [1]. 
2.12.1. SCREEN: Window Width Mode is being selected.
2.13. Select the adaptive width mode to add individual spectra to the averaged spectrum until the maximum sigma target is reached [1]. Allow the recording to stop at the upper limit of the regression window if the sigma target is not reached [2]. Display the measurement results on screen after the minimum sigma target is reached [3].
2.13.1. SCREEN: Adaptive width mode being selected.
2.13.2. SCREEN: Regression window progressing toward the upper limit during data acquisition.
2.13.3. SCREEN: Measurement results appearing after the minimum sigma target is reached.
2.14. Select the fixed-width mode to add individual spectra to the averaged spectrum until the target number of spectra set under average length is included [1].
2.14.1. SCREEN: Fixed-width mode is being selected. Spectra are being accumulated until the target number is reached.
2.15. Once the system settings are complete, click the Plot tab [1]. Under Main Series, select the components to be plotted and displayed on screen during the measurements [2]. Confirm that all selected measurement components are consistent with the protocol of interest [1].
2.15.1. SCREEN: Plot tab being selected.
2.15.2. SCREEN: Main Series section being opened. Measurement components being selected for display.
2.15.3. SCREEN: Selected measurement components being reviewed and confirmed.

3. Data Acquisition
Demonstrator: Click here to enter name of demonstrator(s) 

3.1. Select the appropriate pre-saved settings, recording parameters, and library files [1].
3.1.1. SCREEN: Pre-saved settings, recording parameters, and library files being selected.

3.2. Place the laser probe directly above the tissue or organ of interest [1]. Position the probe as perpendicular as possible and 1 centimeter away from the tissue [2]. Using a ruler, measure 1 centimeter from the surface of the organ [3]. Adjust the probe to a distance less than 1 centimeter if the signal is unclear, while ensuring the probe does not touch the tissue [4].
3.2.1. Talent positioning the laser probe directly above the tissue or organ.
3.2.2. Talent adjusting the laser probe to a perpendicular position 1 centimeter from the tissue.
3.2.3. Talent measuring 1 centimeter from the organ surface with a ruler.
3.2.4. Talent adjusting the probe position closer to the tissue without making contact.
3.3. Confirm that the laser spot is located on the left lateral lobe as close to the portal vein as possible [1].	Comment by Pallavi  Sharma: Authors: Do you want to film this? If yes, please specify the organ used in step 3.2. Otherwise it will confuse the viewers
3.3.1. Shot of the laser spot positioned on the left lateral lobe near the portal vein.
3.4. Confirm that all selected libraries, regression settings, target sigma, regression window, and acquisition modes are correct and as desired [1].
3.4.1. SCREEN: Libraries, regression settings, target sigma, regression window, and acquisition modes being reviewed and confirmed.
3.5. On the main page, click the Start button to begin data acquisition [1]. 
3.5.1. SCREEN: Start button being selected on the main page.
3.6. Once data acquisition and the experimental protocol are complete, discard the entire liver or tissue into a biohazard bin [1]. Clean any system or surface that contacted the organ or tissue with 70% ethanol [2].
3.6.1. Talent discarding the whole liver or tissue into a biohazard bin.
3.6.2. Talent cleaning the contacted system and surfaces with 70% ethanol.
3.7. If working with a perfusion system, run deionized water mixed with an enzymatic cleanser. Then, run clean deionized water through the system [1].
3.7.1. Talent running the enzymatic cleanser solution through the perfusion system.
3.8. Infer the quality of the RRS signal from the plots and values displayed by the software [1]. If the RRS signal does not meet the quality control criteria, adjust the probe to be as perpendicular as possible and approximately 9 millimeters from the surface to the probe window [2].
3.8.1. SCREEN: Plots and values being reviewed to assess the RRS signal quality.
3.8.2. Talent adjusting the probe position relative to the tissue surface.
3.9. Make sure the room is as dark as possible [1]. Verify that the probe window is free of contamination, especially blood [2]. Using a lint-free, non-abrasive wipe, clean the quartz window with 70% isopropyl alcohol if contamination is present [3], and dry it thoroughly with a dry wipe before proceeding [4].
3.9.1. Shot of the darkened experimental room.
3.9.2. Talent inspecting the probe window for contamination.
3.9.3. Talent cleaning the quartz window with a lint-free, non-abrasive wipe and 70% isopropyl alcohol.
3.9.4. Talent drying the quartz window with a dry wipe.
3.10. In case of heavy contamination, wipe the probe thoroughly as shown earlier [1]. Then, place the laser in a dark enclosure lined with black tape [2], press the Take Dark button to calibrate the device, and allow the calibration process to proceed for approximately 5 to 10 minutes [3]. Regardless of contamination, calibrate the device monthly and confirm that automatic reminder messages remain enabled [4].
3.10.1. Talent thoroughly wiping the probe.
3.10.2. Talent placing the laser in a dark enclosure lined with black tape.
3.10.3. SCREEN: Take Dark button being selected.
3.10.4. SCREEN: Automatic calibration reminder messages being displayed in the software.

3.11. Afterward, subject the acquired rat liver to varying levels of ischemia and oxygen reperfusion as a stress test [1]. Take RRS measurements throughout the stress test, take inflow and outflow measurements, and plot the 3RNR values [2].
3.11.1. Representative shot of the liver undergoing stress test.
3.11.2. SCREEN: The RRS values are being recorded.

3.11.3. 

Results
Please review this section to make sure that it accurately reflects your findings.
· This section will not be recorded by the videographer. It only includes the figures/tables from your manuscript (called LAB MEDIA). 
· Use Track Changes when making edits or revisions. Ensure the voiceover length is below 200 words. Current word count: 150.
· Please note that the video cannot include voiceover without an accompanying visual.

4. Results 

4.1. The 3RMR values at T0 (T-Zero) and T5 (T-Five) were within the optimal range of 10% to 30% in continuously oxygenated livers [1].
4.1.1. LAB MEDIA: Figure 5A. Video editor: Highlight the data point between T0 and T5 
4.2. The 3RMR value at the end of the ischemic period, T10, was outside the optimal range [1]. At the end of the reperfusion period, T15, the 3RMR value returned to the optimal range [2].
4.2.1. LAB MEDIA: Figure 5A. Video editor: Highlight the data point at T10 above the shaded optimal range
4.2.2. LAB MEDIA: Figure 5A. Video editor: Highlight the data point at T15 within the shaded optimal range
4.3. Raw RRS spectra at T0 [1] and T15 [2] showed spectral characteristics of oxidized mitochondria, including a spectral marker peak at 1371 inverse centimeters [3].
4.3.1. LAB MEDIA: Figure 6A. 
4.3.2. LAB MEDIA: Figure 6C. 
4.3.3. LAB MEDIA: Figure 6D and F. Video editor: Highlight the marked peak 
4.4. The RRS spectra at T10 showed a reduced marker peak at 1357 inverse centimeters [1].
4.4.1. LAB MEDIA: Figure 6B and E
4.5. An increase in 3RMR levels was observed for complex III and complex IV during ischemia, and both returned to optimal levels after oxygen was reintroduced [1].
4.5.1. LAB MEDIA: Figure 5B. 
4.6. Outflow lactate levels [1] and potassium levels remained relatively consistent throughout the stress test despite changes in oxygenation [2].
4.6.1. LAB MEDIA: Figure 5D. 
4.6.2. LAB MEDIA: Figure 5C. 
4.7. Oxygen consumption rates showed no signs of damage even after a period of no-flow ischemia [1].
4.7.1. LAB MEDIA: Figure 5E. 
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