# Submission ID #: 71290

# Scriptwriter Name: Pallavi Sharma

# Project Page Link: https://review.jove.com/files_upload.php?src=21743498

# Title: Real-Time, Noninvasive Evaluation of Mitochondrial Function Using Resonance Raman Spectroscopy

# Authors and Affiliations:

Yasemin Altug1,2, Rohil Jain1,2, Khanh T. Nguyen1,2, Padraic Romfh3, Shannon N. Tessier1,2

1Center of Engineering in Medicine and Surgery, Department of Surgery, Massachusetts General Hospital, Harvard Medical School

2Shriners Hospital for Children

3Pendar Technologies

# Corresponding Authors:

Shannon N. Tessier 	sntessier@mgh.harvard.edu

# Email Addresses for All Authors:

Yasemin Altug 	yaltug@mgh.harvard.edu

Rohil Jain 		rohilsjain@gmail.com

Khanh T. Nguyen 	ktnguyen360@gmail.com

Padraic Romfh 	promfh@pendar.com

Shannon N. Tessier 	sntessier@mgh.harvard.edu

## Author Questionnaire

1. Microscopy: Does your protocol require the use of a dissecting or stereomicroscope for performing a complex dissection, microinjection technique, or something similar?  NO

2. Software: Does the part of your protocol being filmed include step-by-step descriptions of software usage?  YES

Authors: Please create screen capture videos of the shots labeled as SCREEN, create a screenshot summary, and upload the files to your project page as soon as possible:  https://review.jove.com/files_upload.php?src=21743498

3. Filming location: Will the filming need to take place in multiple locations?   NO

4. Testimonials (optional): Would you be open to filming two short testimonial statements live during your JoVE shoot? These will not appear in your JoVE video but may be used in JoVE’s promotional materials. NO

Current Protocol Length

Number of Steps:  26

Number of Shots:  50

# Interviews

Videographer: Obtain headshots for all authors available at the filming location.

Authors: Please note that the questions will not appear on screen. Please answer in stand-alone full sentence with sufficient context

INTRODUCTION:

REQUIRED: Summarize your research focus and the main questions it aims to answer.

  - Yasemin Altug: My research focuses on utilizing mitochondrial redox states and mitochondrial health as indicators of organ viability during cold-storage and machine perfusion.

    - INTERVIEW: Named Talent says the statement above in an interview-style shot, looking slightly off-camera.

Describe the research gap or limitation in existing methods and how this protocol addresses them.

  - Yasemin Altug: Current methods inadequately assess organ viability during cold preservation, whereas mitochondrial redox measurements provide a temperature-independent indicator of viability.

    - INTERVIEW: Named Talent says the statement above in an interview-style shot, looking slightly off-camera. Suggested B roll: 3.5.1

CONCLUSION:

What does this protocol allow researchers to study or measure?

  - Yasemin Altug: This protocol measures mitochondrial redox states, which in turn allows the users to measure organ viability.

    - INTERVIEW: Named Talent says the statement above in an interview-style shot, looking slightly off-camera.

Describe the most important consideration or challenge when performing this protocol.

  - Yasemin Altug: A critical consideration while performing this protocol is the creation and validation of tissue- and wavelength-specific spectral libraries, as new libraries may be required for each application.

    - INTERVIEW: Named Talent says the statement above in an interview-style shot, looking slightly off-camera.

Are there any additional methods or analyses that can be performed following this procedure?

  - Yasemin Altug: Additional analyses that can be performed using this procedure include ATP quantification, inflammatory biomarkers, and liver injury markers.

    - INTERVIEW: Named Talent says the statement above in an interview-style shot, looking slightly off-camera.

Videographer: Obtain headshots for all authors available at the filming location.

# Ethics Title Card

This research has been approved by the Institutional Animal Care and Use Committee (IACUC) at Massachusetts General Hospital

# Protocol

Videographer: Please always record an audio slate while filming each shot mentioned in the shotlist.

- Setting up the Resonance Raman Spectroscopy (RRS) System

Demonstrator: Yasemin Altug

  - To begin, connect the spectrometer unit to the direct current power supply [1] and connect it to the universal serial bus port on the computer running the analysis software using the supplied cables [2]. Identify the spectrometer unit with the round plug that contains alignment indentations [3].

    - WIDE: Talent connecting the spectrometer unit to the direct current power supply. Videographer: Please obtain CU shot as well for this shot.

    - Talent connecting the spectrometer unit to the universal serial bus port on the computer using the supplied cable.

    - Close-up of the spectrometer unit showing the round plug with alignment indentations.

  - Align the indentations of the round plug with those of the cable connected to the power supply and universal serial bus port [1]. Connect the ports without cleaning them beforehand [2].

    - Talent aligning the indentations of the round plug with the cable connector.

    - Talent connecting the ports together.

  - Place the laser probe in the probe holder [1]. Then adjust the probe's position at the start of the experiment [2].

    - Talent placing the laser probe in the probe holder.

    - Talent adjusting the position of the laser probe.

  - Open the spectrometer control software [1]. Observe the graph showing the percent Resonance Raman reduced mitochondrial ratio or 3RMR (Three-R-M-R), tissue oxygen saturation, and myoglobin saturation upon opening the spectrometer control software [2]. Click the Settings tab at the bottom left of the page [3].

    - SCREEN: Spectrometer control software being opened.

Authors: Please create screen capture videos of the shots labeled as SCREEN, create a screenshot summary, and upload the files to your project page as soon as possible:  https://review.jove.com/files_upload.php?src=21743498

    - SCREEN: Graph displaying the percent 3RMR, tissue oxygen saturation, and myoglobin saturation metrics in the spectrometer control software.

    - SCREEN: Settings tab being selected at the bottom left of the page.

  - Under the Info tab, enter the session information in Session Name [1] and any necessary details in Session Description [2].

    - SCREEN: Info tab being opened. Session information being entered into Session Name.

    - SCREEN: Details being entered into Session Description.

  - Select Continuous Mode as the data collection mode for the experiment. Confirm that spectra will be recorded continuously until the stop button is pressed [2].

    - SCREEN: Continuous Mode being selected as the data collection mode. Continuous recording mode being displayed in the software.

  - Once all session information is entered, click the System tab [1].

    - SCREEN: System tab being selected.

  - Under Settings File, select the protocol appropriate for the project [1].

    - SCREEN: Settings File menu being opened and a protocol being selected.

  - To create a new settings file, adjust the required settings parameters. Then, click the Save button to save the parameters as a separate settings file [1].

    - SCREEN: Settings parameters being modified. Save button being selected and the settings file being saved.

  - To import a new settings file, click the drop-down menu next to Default (Mod) (Default -Mod) [1]. Select Import New Settings File and click the new settings file to import it into the system [2].

    - SCREEN: Drop-down menu next to Default (Mod) being opened.

    - SCREEN: Import New Settings File being selected. New settings file being selected and imported into the system.

  - In the Regression Libraries section, select and confirm the libraries for data acquisition and analysis. The libraries for reduced and oxidized mitochondria, complex III, complex IV, hemoglobin, and myoglobin were recorded before the experiment [1].

    - SCREEN: Regression Libraries section being opened. The library combination is being selected for data acquisition and analysis.

  - Select and confirm the preferred Window Width Mode [1].

    - SCREEN: Window Width Mode is being selected.

  - Select the adaptive width mode to add individual spectra to the averaged spectrum until the maximum sigma target is reached [1]. Allow the recording to stop at the upper limit of the regression window if the sigma target is not reached [2]. Display the measurement results on screen after the minimum sigma target is reached [3].

    - SCREEN: Adaptive width mode being selected.

    - SCREEN: Regression window progressing toward the upper limit during data acquisition.

    - SCREEN: Measurement results appearing after the minimum sigma target is reached.

  - Select the fixed-width mode to add individual spectra to the averaged spectrum until the target number of spectra set under average length is included [1].

    - SCREEN: Fixed-width mode is being selected. Spectra are being accumulated until the target number is reached.

  - Once the system settings are complete, click the Plot tab [1]. Under Main Series, select the components to be plotted and displayed on screen during the measurements [2]. Confirm that all selected measurement components are consistent with the protocol of interest [1].

    - SCREEN: Plot tab being selected.

    - SCREEN: Main Series section being opened. Measurement components being selected for display.

    - SCREEN: Selected measurement components being reviewed and confirmed.

- Data Acquisition

  - Select the appropriate pre-saved settings, recording parameters, and library files [1].

    - SCREEN: Pre-saved settings, recording parameters, and library files being selected.

  - Place the laser probe directly above the rat liver or any other organ of interest [1]. Position the probe as perpendicular as possible and 1 centimeter away from the tissue [2]. Using a ruler, measure 1 centimeter from the surface of the organ [3]. Adjust the probe to a distance less than 1 centimeter if the signal is unclear, while ensuring the probe does not touch the tissue [4].

    - Talent positioning the laser probe directly above the tissue or organ.

    - Talent adjusting the laser probe to a perpendicular position 1 centimeter from the tissue.

    - Talent measuring 1 centimeter from the organ surface with a ruler.

    - Talent adjusting the probe position closer to the tissue without making contact.

  - Confirm that the laser spot is located on the left lateral lobe as close to the portal vein as possible [1].

    - Shot of the laser spot positioned on the left lateral lobe near the portal vein.

  - Confirm that all selected libraries, regression settings, target sigma, regression window, and acquisition modes are correct and as desired [1].

    - SCREEN: Libraries, regression settings, target sigma, regression window, and acquisition modes being reviewed and confirmed.

  - On the main page, click the Start button to begin data acquisition [1].

    - SCREEN: Start button being selected on the main page.

  - Once data acquisition and the experimental protocol are complete, discard the entire liver or tissue into a biohazard bin [1]. Clean any system or surface that came into contact with the organ or tissue using 70% ethanol [2].

    - Talent discarding the whole liver or tissue into a biohazard bin.

    - Talent cleaning the contacted system and surfaces with 70% ethanol.

  - If working with a perfusion system, run deionized water mixed with an enzymatic cleanser. Then, run clean deionized water through the system [1].

    - Talent running the enzymatic cleanser solution through the perfusion system.

  - Infer the quality of the RRS signal from the plots and values displayed by the software [1]. If the RRS signal does not meet the quality control criteria, adjust the probe to be as perpendicular as possible and approximately 9 millimeters from the surface to the probe window [2].

    - SCREEN: Plots and values being reviewed to assess the RRS signal quality.

    - Talent adjusting the probe position relative to the tissue surface.

  - Make sure the room is as dark as possible [1]. Verify that the probe window is free of contamination, especially blood [2]. Using a lint-free, non-abrasive wipe, clean the quartz window with 70% isopropyl alcohol if contamination is present [3], and dry it thoroughly with a dry wipe before proceeding [4].

    - Shot of the darkened experimental room.

    - Talent inspecting the probe window for contamination.

    - Talent cleaning the quartz window with a lint-free, non-abrasive wipe and 70% isopropyl alcohol.

    - Talent drying the quartz window with a dry wipe.

  - In case of heavy contamination, wipe the probe thoroughly as shown earlier [1]. Then, place the laser in a dark enclosure lined with black tape [2], press the Take Dark button to calibrate the device, and allow the calibration process to proceed for approximately 5 to 10 minutes [3]. Regardless of contamination, calibrate the device monthly and confirm that automatic reminder messages remain enabled [4].

    - Talent thoroughly wiping the probe.

    - Talent placing the laser in a dark enclosure lined with black tape.

    - SCREEN: Take Dark button being selected.

    - SCREEN: Automatic calibration reminder messages being displayed in the software.

  - Afterward, subject the acquired rat liver to varying levels of ischemia and oxygen reperfusion as a stress test [1]. Take RRS measurements throughout the stress test, take inflow and outflow measurements, and plot the 3RNR values [2].

    - Representative shot of the liver undergoing stress test.

    - SCREEN: The RRS values are being recorded.

# Results

- Results

  - The 3RMR values at T0 (T-Zero) and T5 (T-Five) were within the optimal range of 10% to 30% in continuously oxygenated livers [1].

    - LAB MEDIA: Figure 5A. Video editor: Highlight the data point between T0 and T5

  - The 3RMR value at the end of the ischemic period, T10, was outside the optimal range [1]. At the end of the reperfusion period, T15, the 3RMR value returned to the optimal range [2].

    - LAB MEDIA: Figure 5A. Video editor: Highlight the data point at T10 above the shaded optimal range

    - LAB MEDIA: Figure 5A. Video editor: Highlight the data point at T15 within the shaded optimal range

  - Raw RRS spectra at T0 [1] and T15 [2] showed spectral characteristics of oxidized mitochondria, including a spectral marker peak at 1371 inverse centimeters [3].

    - LAB MEDIA: Figure 6A.

    - LAB MEDIA: Figure 6C.

    - LAB MEDIA: Figure 6D and F. Video editor: Highlight the marked peak

  - The RRS spectra at T10 showed a reduced marker peak at 1357 inverse centimeters [1].

    - LAB MEDIA: Figure 6B and E

  - An increase in 3RMR levels was observed for complex III and complex IV during ischemia, and both returned to optimal levels after oxygen was reintroduced [1].

    - LAB MEDIA: Figure 5B.

  - Outflow lactate levels [1] and potassium levels remained relatively consistent throughout the stress test despite changes in oxygenation [2].

    - LAB MEDIA: Figure 5D.

    - LAB MEDIA: Figure 5C.

  - Oxygen consumption rates showed no signs of damage even after a period of no-flow ischemia [1].

    - LAB MEDIA: Figure 5E.