# Submission ID #: 71304

# Scriptwriter Name: Poornima G

# Project Page Link: https://review.jove.com/account/file-uploader?src=21747768

# Title: Multicenter Flow Cytometry for Acute Myeloid Leukemia Measurable Residual Disease with LAIP/DfN and CD34+CD38− Leukemic Stem Cell-enriched Populations

# Authors and Affiliations:

# Adriana Plesa1,35*, Florent Dumezy2, Stéphanie Mathis3, Isabelle Arnoux4, Veronique Saada5  ,Valerie Bardet6 , Iole Oster1 ,Bouchra Badaoui7, Elsa Bera8, Edouard Bonneville9 , Boutheina Bouslama10 ,  Nicolas Chapuis11 ,Clementine Chauvel3 , Edouard Cornet12 , Louis Thomas Dannus13 , Camille Debord14 , Fréderic Feger15 , Jean Feuillard16  ,Thomas Fournet17, Francine Garnache17 , Franck Geneviève18, Delphine Gerard19, Valerie Goncalves Monteiro20, Julien Guy21 ,Veronique Harrivel22, Claire Hemar23, Cassandra Jadeau16, Caroline Mayeur-Rousse14,  Hélène Lapillonne10,  Veronique Latger-Cannard19 , Magali Le Garff-Tavernier24,  Pierre Lemaire3, Remi Letestu25  , Anne Catherine Lhoumeau26, Delphine Manzoni1, Aguirre Mimoun27 ,Vanessa Nivaggioni4,  Jennifer Osman28,  Victoria Raggueneau28 , Hubert Rambaud18, Tatiana Raskovalova9 , Anne Roggy17 , Mikael Roussel29 , Didier Sciortino26 , Valerie Soenen2, François Vergez30 ,Richard Veyrat Masson13, Jean Philippe Vial27 , Manon Vogrig31 , Orianne Wagner-Ballon7,37 , Jaja Zhu6 , Karine Celli-Lebras32, Raphaël Itzykson33,38, Christian Récher34, Hervé Dombret32,33,Claude Preudhomme2;36 , Christophe Roumier2,36*

# * Coordinators of the Acute Leukemia French Intergroup AML MRDflow & LSC network

# ¹Laboratory of Hematology and Flow cytometry, Lyon-Sud Hospital, Hospices Civils de Lyon, CHU-Lyon

# ²Laboratory of Hematology and Flow cytometry, CHU-Lille

# ³Laboratory of Hematology and Flow cytometry, Saint Louis Hospital, APHP

# ⁴Laboratory of Hematology and Flow cytometry, APHM-Marseille

# ⁵Laboratory of Hematology and Flow cytometry, IGR

# ⁶Laboratory of Hematology and Flow cytometry, Ambroise Pare University Hospital, APHP, Paris Saclay

# ⁷Département d’Hématologie et Immunologie Biologiques, Hôpitaux Universitaires Henri Mondor, APHP

# ⁸Laboratory of Hematology and Flow cytometry, CHU-Rouen

# ⁹Laboratory of Immunology and Flow cytometry, CHU-Grenoble

# ¹⁰Laboratory of Hematology and Flow cytometry, Trousseau Hospital, APHP

# ¹¹Laboratory of Hematology and Flow cytometry, Cochin Hospital, APHP

# ¹²Laboratory of Hematology and Flow cytometry, CHU-Caen

# ¹³Laboratory of Hematology and Flow cytometry, CHU-Clermont Ferrand

# ¹⁴Laboratory of Hematology and Flow cytometry, CHU-Nantes

# ¹⁵Laboratory of Hematology and Flow cytometry, Saint Antoine Hospital, APHP

# ¹⁶Laboratory of Hematology and Flow cytometry, CHU-Limoges

# ¹⁷Laboratory of Hematology and Flow cytometry, CHU-Besancon

# ¹⁸Laboratory of Hematology and Flow cytometry, CHU-Angers

# ¹⁹Laboratory of Hematology and Flow cytometry, CHU-Nancy

# ²⁰Laboratory of Hematology and Flow cytometry, CHU-Reims

# ²¹Laboratory of Hematology and Flow cytometry, CHU-Dijon

# ²²Laboratory of Hematology and Flow cytometry, CHU-Amiens

# ²³Laboratory of Hematology and Flow cytometry, CH-Valenciennes

# ²⁴Laboratory of Hematology and Flow cytometry, Pitie Salpetriere Hospital, APHP

# ²⁵Laboratory of Hematology and Flow cytometry, Avicenne Hospital, APHP

# ²⁶Laboratory of Hematology and Flow cytometry, IPC-Marseille

# ²⁷Laboratory of Hematology and Flow cytometry, CHU-Bordeaux

# ²⁸Laboratory of Hematology and Flow cytometry, Versailles Hospital

# ²⁹Laboratory of Hematology and Flow cytometry, CHU-Rennes

# ³⁰Laboratory of Hematology and Flow cytometry, CHU-Toulouse

# ³¹Laboratory of Hematology and Flow cytometry, CHU-Saint Etienne

# ³²Acute Leukemia French Association Group Coordination, IRSL

# ³³Department of Hematology, Saint Louis Hospital, APHP

# ³⁴Department of Hematology, CHU Toulouse, Université de Toulouse, Institut Universitaire du Cancer de Toulouse Oncopole

# ³⁵CRCL INSERM 1052/CNRS 5286, University of Lyon

# ³⁶UMR9020 CNRS-UMR-S1277 INSERM, University of Lille

# ³⁷Université Paris-Est Créteil (UPEC), INSERM U955 IMRB

# ³⁸Service Hématologie Adultes, Hôpital Saint Louis, APHP Paris, Université Paris Cité, INSERM, CNRS, Institut de Recherche Saint Louis (IRSL)

# Corresponding Author:

# Adriana Plesa		adriana.plesa@chu-lyon.fr

# Email Addresses for All Authors:

Florent Dumezy 	 florent.dumezy@chu-lille.fr

Stéphanie Mathis 	stephanie.mathis@aphp.fr

Isabelle Arnoux 	 isabelle.arnoux@mail.ap-hm.fr

Veronique Saada 	veronique.saada@igr.fr

Valerie Bardet 	 valerie.bardet@aphp.fr

Iole Oster 	 	iole.oster@chu-lyon.fr

Bouchra Badaoui 	 bouchra.badaoui@aphp.fr

Elsa Bera 	 	Elsa.Bera@chu-rouen.fr

Edouard Bonneville 	 ebonneville@chu-grenoble.fr

Boutheina Bouslama 	boutheina.bouslama@aphp.fr

Nicolas Chapuis 	nicolas.chapuis@aphp.fr

Clementine Chauvel 	 clementine.chauvel@aphp.fr

Edouard Cornet 	 cornet-e@chu-caen.fr

Louis Thomas Dannus  ltdannus@chu-clermontferrand.fr

Camille Debord 	 camille.debord@chu-nantes.fr

Fréderic Feger 	frederic.feger@aphp.fr

Jean Feuillard 		jean.feuillard@unilim.fr

Thomas Fournet 	 tfournet@chu-besancon.fr

Francine Garnache 	 fgarnache@chu-besancon.fr

Franck Geneviève 	 FrGenevieve@chu-angers.fr

Delphine Gerard 	 D.GERARD@chru-nancy.fr

Valerie Goncalves Monteiro  vgoncalves-monteiro@chu-reims.fr

Julien Guy 		julien.guy@chu-dijon.fr

Veronique Harrivel 	 Harrivel.Veronique@chu-amiens.fr

Claire Hemar		 hemar-c@ch-valenciennes.fr

Cassandra Jadeau 	Cassandra.JADEAU@chu-limoges.fr

Caroline Mayeur-Rousse	 caroline.mayeurrousse@chu-nantes.fr

Hélène Lapillonne 	helene.lapillonne@aphp.fr

Veronique Latger-Cannard 	v.cannard@chru-nancy.fr

Magali Le Garff-Tavernier 	 magali.legarff@aphp.fr

Pierre Lemaire 	pierre.lemaire@aphp.fr

Remi Letestu 		remi.letestu@aphp.fr

Anne Catherine Lhoumeau 	LHOUMEAUAC@ipc.unicancer.fr

Delphine Manzoni 	delphine.manzoni@chu-lyon.fr

Aguirre Mimoun 	aguirre.mimoun@chu-bordeaux.fr

Vanessa Nivaggioni	 Vanessa.NIVAGGIONI@ap-hm.fr

Jennifer Osman 	josman@ch-versailles.fr

Victoria Raggueneau 	vraggueneau@ch-versailles.fr

Hubert Rambaud	 HuRambaud@chu-angers.fr

Tatiana Raskovalova 	TRaskovalova@chu-grenoble.fr

Anne Roggy		 aroggy@chu-besancon.fr

Mikael Roussel 	Mikael.ROUSSEL@chu-rennes.fr

Didier Sciortino 	 SCIORTINOD@ipc.unicancer.fr

Valerie Soenen 	Valerie.SOENEN@ chu-lille.fr

François Vergez 	Vergez.Francois@iuct-oncopole.fr

Richard Veyrat Masson 	rveyrat@chu-clermontferrand.fr

Jean Philippe Vial 	jean-philippe.vial@chu-bordeaux.fr

Manon Vogrig		 Manon.VOGRIG@chu-st-etienne.fr

Orianne Wagner-Ballon 	 orianne.wagnerballon@aphp.fr

Jaja Zhu 		jaja.zhu@aphp.fr

Karine Celli-Lebras 	karine.cellilebras@alfa-leukemia.org

Raphaël Itzykson 	raphael.itzykson@aphp.fr

Christian Recher	 Recher.Christian@iuct-oncopole.fr

Hervé Dombret	 herve.dombret@mac.com

Claude Preudhomme	Claude.PREUDHOMME@CHU-LILLE.FR

Christophe Roumier	Christophe.ROUMIER@chu-lille.fr

Adriana Plesa		adriana.plesa@chu-lyon.fr

## Author Questionnaire

1. Microscopy: Does your protocol require the use of a dissecting or stereomicroscope for performing a complex dissection, microinjection technique, or something similar?  YES

2. Software: Does the part of your protocol being filmed include step-by-step descriptions of software usage?  YES, all done

3. Filming location: Will the filming need to take place in multiple locations?   NO

4. Testimonials (optional): Would you be open to filming two short testimonial statements live during your JoVE shoot? These will not appear in your JoVE video but may be used in JoVE’s promotional materials. NO

Current Protocol Length

Number of Steps:  25

Number of Shots:  55 (35 SC)

# Interviews

Videographer: Obtain headshots for all authors available at the filming location.

Authors: Please note that the questions will not appear on screen. Please answer in stand-alone full sentence with sufficient context

We have retained 5 best answers as per our journal guidelines

INTRODUCTION:

REQUIRED: Summarize your research focus and the main questions it aims to answer.

  - Adriana Plesa: Our research focused on accurate quantification of measurable residual disease, including Leukemic Stem Cells, in acute myeloid leukemia patients.

    - INTERVIEW: Named talent says the statement above in an interview-style shot, looking slightly off-camera

Describe the research gap or limitation in existing methods and how this protocol addresses them.

  - Christophe Roumier: To implement multicenter and multiplatform monitoring, we tailored the flow cytometer settings to each platform and a common analysis strategy.

    - INTERVIEW: Named talent says the statement above in an interview-style shot, looking slightly off-camera

CONCLUSION:

What does this protocol allow researchers to study or measure?

  - Iole Oster: This protocol allows researchers to measure the Leukemia aberrant immunophenotype identified at diagnosis and leukemic stem cells in the CD34+CD38- enriched fraction.

    - INTERVIEW: Named talent says the statement above in an interview-style shot, looking slightly off-camera

Describe the most important consideration or challenge when performing this protocol.

  - Herve Dombret: Choosing a multicenter approach can be challenging. The feasibility of this concept requires mandatory steps before implementing MRD flow in AML clinical trials. (take by SCREEN AP/CR)

    - INTERVIEW: Named talent says the statement above in an interview-style shot, looking slightly off-camera

How can future studies build on this work? Describe the questions or research directions that can be investigated next.

  - Adriana Plesa: The "crossroads" moment has arrived to move forward and incorporate LSC quantification performed with great accuracy into future AML clinical trials at diagnosis and follow-up.

    - INTERVIEW: Named talent says the statement above in an interview-style shot, looking slightly off-camera

# Videographer: Obtain headshots for all authors available at the filming location.

# Ethics Title Card

This research has been approved by the Human Ethics Committee of CHU Lyon and the Ethics Committee of the Institutional Review Board of Lille CHU

# Protocol

Videographer: Please always record an audio slate while filming each shot mentioned in the shotlist.

- Cytometer Standardization

Demonstrator:  Iole Oster

  - To begin, prepare the 8-peak rainbow beads and determine the cytometer-specific target values [1].

    - WIDE: Talent preparing the 8-peak rainbow beads. Videographer: Please obtain CU shot as well for this shot.

  - Set each machine according to the target values [1] and ensure that each channel's measured value is within 5% of its target [2].

    - Talent setting the machine according to the target values.

    - LAB MEDIA: Table 1.

NOTE: Treat this as separate section

Bone Marrow Sample Preparation

  - Collect 1 to 2 milliliters of bone marrow in EDTA from acute myeloid leukemia patients at diagnosis, during treatment monitoring, and at relapse, when applicable [1-TXT]. For morphologic examination, place a drop of bone marrow aspirate on a glass slide [2] and gently spread the sample [3]. After 30 minutes, stain the smear with May-Grunewald-Giemsa [4] and examine the sample under a microscope at 100x magnification [5]. Next, add 15 milliliters of ammonium chloride lysing solution to the collected bone marrow sample [6]. Mix gently and incubate for 10 minutes at room temperature [7].

    - Talent adding bone marrow sample to EDTA tubes and labeling the tubes as ‘Diagnosis’, MRD-1. TXT: Treatment monitoring: MRD-1 to 4; MRD: Measurable Residual Disease  Authors: This labeling is just for representation

    - Talent adding a drop of sample on a slide.

    - Talent spreading then sample.

    - Talent adding the stain on the sample.

    - LAB MEDIA: Supplementary figure 10. Video editor: Label the left panel as ‘Healthy regenerating marrow’ and right one as “Acute myeloid leukemia patient”

    - Talent adding ammonium chloride lysing solution to the bone marrow sample.

    - Talent gently mixing the sample and keeping it aside.

  - Then, centrifuge the sample at 450 g for 5 minutes at room temperature [1] and discard the supernatant [2]. Resuspend the cell pellet in 10 milliliters of PBS with 0.1% BSA [3]. Repeat the centrifugation step at 450 g for 5 minutes [4]. After discarding the supernatant, resuspend the cell pellet at a concentration of 30 × 10⁶ white blood cells per milliliter in PBS [6].

    - Talent placing the sample in the centrifuge.

    - Talent discarding the supernatant.

    - Talent resuspending the cell pellet in PBS with 0.1% bovine serum albumin.

    - Talent placing the sample in the centrifuge.

    - Talent resuspending the cell pellet in PBS.

- Immunostaining the Bone Marrow Cells

  - Distribute the monoclonal antibodies in tubes T1, T2, and T3 according to the platform [1-TXT].

    - Talent distributing monoclonal antibodies into tubes T1, T2, and T3. TXT: See Table 2 for antibody panel configurations

  - Add 2 to 3 × 10⁶ lysed cells per tube in 100 microliters of cell suspension [1] and incubate the cells with antibodies for 15 minutes in the dark at room temperature [2].

    - Talent adding lysed cells to each tube.

    - Talent placing the tubes in a dark chamber for incubation at room temperature.

  - Wash the cells by adding 3 milliliters of PBS [1] and centrifuging at 450 g for 5 minutes [2]. Discard the supernatant [3] and resuspend the stained cells in 300 microliters of PBS [4].

    - Talent adding PBS to the cells.

    - Talent placing the tubes in the centrifuge.

    - Talent discarding the supernatant.

    - Talent resuspending the stained cells in PBS.

- Flow Cytometry

  - Open the MRD assay or acquisition protocol with the adapted settings and compensation matrix [1].

    - SCREEN: MRD assay or acquisition protocol being opened with the adapted settings and compensation matrix.  NOTE: Use videographer’s footage for 4.1.1 and 4.2.2

Videographer: Please record the computer screen for the shots labeled as SCREEN as back-up. The screen should be in focus and the fonts should be clearly visible.

  - Load the samples and acquire a minimum of 500,000 to 1,000,000 white blood cells [1] in each tube at diagnosis and follow-up [2].

    - Talent loading the samples.

    - SCREEN: Acquisition in process.

NOTE: Treat this as a separate section

Leukemia-Associated Immunophenotype (LAIP) Analysis

  - Use a common template and dot-plot sequence [1].

    - SCREEN: SCREEN-4.3 00:00-00:05

  - Capture leukemia-associated immunophenotype, patterns in bulk blasts by drawing a single-cell gate based on forward scatter width versus forward scatter height [1].

    - SCREEN: SCREEN-4.4 00:05-00:21

  - Draw a live-cell gate based on forward scatter versus side scatter [1].

    - SCREEN: SCREEN-4.5 00:07-00:16

  - Then, draw a blast gate on CD45 dim and side scatter [1], followed by the primitive marker gate using CD34 and CD117 [2].

    - SCREEN: SCREEN-4.6.1 00:05-00:17

    - SCREEN: SCREEN-4.6.2 00:05-00:17

  - Using fixed biparametric dot plots, draw the relevant gates, including LAIP 1 through LAIP 8 [1]. After identifying and excluding relevant physiological and nonspecific populations, apply the Boolean gating strategy for LAIP analysis [2-TXT].

    - SCREEN: SCREEN-4.7.1 00:06-00:11 AND 00:35-00:38

    - SCREEN:  SCREEN-4.7.2 00:10-00:22. TXT: See text protocol for the LAIP approach Boolean operators

- Different-from-Normal (DfN) Analysis

  - To capture a DfN (D-F-N) pattern in bulk blasts, draw live cells using forward scatter and side scatter [2].

    - SCREEN: SCREEN-5.1.2 00:05-00:14

  - Draw single cells by removing doublets based on the forward scatter peak and the side scatter area [1]. Then, draw the mononuclear cell gate using forward scatter and side scatter [2].

    - SCREEN: SCREEN-5.2.1 00:07-00:20

    - SCREEN: SCREEN-5.2.2 00:08-00:18

  - Apply a cleanup gate to remove lymphocytes using CD45, CD33, and side scatter [1]. Next, apply a cleanup gate to remove erythroblasts and debris using CD45 and CD36 [2].

    - SCREEN: SCREEN-5.3.1 00:06-00:16

    - SCREEN: SCREEN-5.3.2 00:04-00:06 and 00:13-00:18

  - Use two input gates, CD34 positive, CD45, and side scatter [1], and CD117 positive, CD45, and side scatter, to define the empty boxes [2].

    - SCREEN: SCREEN-5.4.1 00:05-00:10

    - SCREEN: SCREEN-5.4.1 00:13-00:20

  - Draw each DfN empty box based on internal control populations, including lymphocytes for cross-lineage markers and normal myeloid precursors for loss, decrease, or overexpression markers [1]. Compare the empty boxes with reference bone marrow samples, including regenerating, inflammatory, and healthy donor bone marrow. Define abnormal populations as clusters of abnormal cells separated from normal cells by at least 0.5 decade on the log scale [2].

    - SCREEN: SCREEN-5.5.1 00:10-00:25

    - SCREEN: SCREEN-5.5.2 00:04-00:20

- Leukemic Stem Cell Analysis and CD38 Threshold Definition

  - Define the leukemic stem cell, or LSC, input gate as mononuclear cell CD34-positive [1].

    - SCREEN: SCREEN-6.1.1 00:01-00:10

  - Next, set the CD38 channel high enough to distinguish CD38-negative populations from CD38-positive populations with good resolution [1].

    - SCREEN: SCREEN-6.2 00:01-00:10 and 00:18-00:19

  - Use bright reference populations to establish the CD38 signal range [1]. Identify plasma cells as the brightest population and hematogones according to the measurement range of each platform [2].

    - SCREEN: SCREEN-6.3.1 00:00-00:05

    - SCREEN: SCREEN-6.3.2 00:00-00:05 and 00:10-00:12

  - Then, define the CD34-positive CD38-negative population P6 using the CD38 fluorescence-minus-one approach validated by an isoclonic CD38 control [1-TXT]. Outline P7, the CD34-positive and CD38-low population, using the hematogone-negative level [2]. Define P8, the CD34-positive and CD38-high population, using the hematogone-positive and plasma-cell negative levels [3].

    - SCREEN: SCREEN-6.4.1 00:03-00:13 TXT: See text protocol for the Boolean operators

    - SCREEN: SCREEN-6.4.2

    - SCREEN: SCREEN-6.4.3 00:05-00:16

  - Apply the same CD38-negative threshold across the harmonized group, with robust methods such as CD38 fluorescence minus one or an isoclonic CD38 control [1].

    - SCREEN: SCREEN-6.5.1 00:05-00:16

  - Include at least one aberrant immunophenotype in the CD34-positive CD38-negative fraction to define LSC-enrichment [1]. Use LSC markers or lineage markers for this identification [2-TXT]. Use CD36 to remove nonspecific events from the final MRD LSC gate [3].

    - SCREEN: SCREEN-6.6.1 00:03-00:08 and 00:11-00:16

    - SCREEN: SCREEN-6.6.2 00:05-00:13. TXT: LSC markers: CLL1/TIM3/CD97/CD45RA/CD123

Lineage markers: CD7/CD19/CD56/CD33

    - SCREEN: SCREEN-6.1.3 00:03-00:06 and 00:10-00:15

  - Apply the assay sensitivity thresholds, with a limit of quantification of 10⁻⁴ and a limit of detection of 10⁻⁵ [1]. Finally, define LSC positivity at diagnosis as at least 1% of CD45- side-scatter-positive blasts [2], and at MRD follow-up as at least 0.01% of white blood cells. Provide the harmonized clinical flow report [3].

    - SCREEN: SCREEN-6.7.1 00:05-00:17

    - SCREEN: SCREEN-6.7.2 00:05-00:14

    - SCREEN: SCREEN-6.7.2 00:16-00:25

# Results

- Results

  - After applying the defined settings, the cell staining patterns overlapped across the 4 flow cytometer platforms [1], with the main difference being a translation of the cell clouds along the logarithmic scales [2].

    - LAB MEDIA: Figure 6.

    - LAB MEDIA: Figure 6. Video editor: Emphasize that the cell clouds (with red centers) in 2nd row as an example.

  - The wet external quality assessment or EQA showed reproducible distributions of P6 cells [1] and mononuclear cell CD34 positive cells among the white blood CD45 positive cells across the 5 assessment rounds [2].

    - LAB MEDIA: Figure 7A.

    - LAB MEDIA: Figure 7B.

  - More than 90% of the locally analyzed data were included within the centralized interval based on the series mean in Dry-EQA-2017 (20-17) [1] and Dry-EQA-2024 (20-24) [2].

    - LAB MEDIA: Figure 8A. Video editor: Highlight the data points clustered around the central reference line

    - LAB MEDIA: Figure 8B. Video editor: Highlight the data points clustered around the central reference line

  - The leukemia-associated immunophenotype cartography established at diagnosis [1] was applied to the follow-up sample to quantify measurable residual disease [2].

    - LAB MEDIA: Figure 9A. Video editor: Highlight the black cell population marked as the leukemia-associated immunophenotype in the diagnosis panels.

    - LAB MEDIA: Figure 9A. Video editor: Highlight the black measurable residual disease population in the follow-up panels.

  - The DfN approach identified abnormal cells in the areas not observed in normal bone marrow [1].

    - LAB MEDIA: Figure 9B. Video editor: Highlight the black potted regions.

  - The leukemic stem cell-enriched cells expressing the defined marker profile were identified at diagnosis within the CD34-positive CD38-negative fraction [1], and the measurable residual disease leukemic stem cell-enriched cells with the same profile persisted at follow-up [2].

    - LAB MEDIA: Figure 10A. Video editor: Highlight the orange cell population in the diagnosis samples.

    - LAB MEDIA: Figure 10B. Video editor: Highlight the orange cell population in the follow-up samples.

  - The limit of blank for leukemic stem cell detection was less than 0.0005% [1], corresponding to 14 events in the leukemic stem cell Boolean gate [2].

    - LAB MEDIA: Figure 11A. Video editor: Highlight the value showing  0.0003%.

    - LAB MEDIA: Figure 11A. Video editor: Highlight the boxed cluster labelled with “14” events on the right.

  - The serial dilution of leukemic stem cell-enriched cells showed a correlation between the measured measurable residual disease values and the theoretical values [1].

    - LAB MEDIA: Figure 11B. Video editor: Highlight the points aligned along the correlation line in the graph.

  - The serial dilution of blasts with different leukemia-associated immunophenotypes showed a correlation between the measured measurable residual disease values and the theoretical values [1].

    - LAB MEDIA: Figure 12A. Video editor: Highlight the data points following the trend line in the dilution graphs.

  - The limit of quantification reached the 0.1% clinical threshold for all tested leukemia-associated immunophenotypes [1].

    - LAB MEDIA: Figure 12B. Video editor: Highlight the blue bars and the red horizontal line marking the 0.1% clinical threshold.

- Cytometer
Pronunciation link: https://www.merriam-webster.com/medical/cytometer
IPA: /saɪˈtɑːmətər/
Phonetic Spelling: sy·taa·muh·ter

- Myeloid
Pronunciation link: https://www.merriam-webster.com/dictionary/myeloid
IPA: /ˈmaɪəˌlɔɪd/
Phonetic Spelling: my·uh·loyd

- Leukemia
Pronunciation link: https://www.merriam-webster.com/dictionary/leukemia
IPA: /luːˈkiːmiə/
Phonetic Spelling: loo·kee·mee·uh

- Morphologic
Pronunciation link: https://www.merriam-webster.com/dictionary/morphologic
IPA: /ˌmɔr.fəˈlɑː.dʒɪk/
Phonetic Spelling: mor·fuh·laa·jik

- Aspirate
Pronunciation link: https://www.merriam-webster.com/dictionary/aspirate
IPA: /ˈæs.pə.rət/
Phonetic Spelling: as·puh·ruht

- Centrifuge
Pronunciation link: https://www.merriam-webster.com/dictionary/centrifuge
IPA: /ˈsɛn.trəˌfjuːdʒ/
Phonetic Spelling: sen·truh·fyooj

- Supernatant
Pronunciation link: https://www.merriam-webster.com/dictionary/supernatant
IPA: /ˌsuː.pərˈneɪ.tənt/
Phonetic Spelling: soo·per·nay·tuhnt

- Monoclonal
Pronunciation link: https://www.merriam-webster.com/dictionary/monoclonal
IPA: /ˌmɑː.nəˈkloʊ.nəl/
Phonetic Spelling: maa·nuh·kloh·nuhl

- Immunophenotype
Pronunciation link: https://www.merriam-webster.com/medical/immunophenotype
IPA: /ˌɪm.jə.noʊˈfiː.nəˌtaɪp/
Phonetic Spelling: im·yuh·noh·fee·nuh·type

- Biparametric
Pronunciation link: No confirmed link found
IPA: /ˌbaɪˌpærəˈmɛtrɪk/
Phonetic Spelling: by·pair·uh·meh·trik

- Lymphocytes
Pronunciation link: https://www.merriam-webster.com/dictionary/lymphocyte
IPA: /ˈlɪm.fəˌsaɪts/
Phonetic Spelling: lim·fuh·sites

- Erythroblasts
Pronunciation link: https://www.merriam-webster.com/dictionary/erythroblast
IPA: /ɪˈrɪθ.rəˌblæsts/
Phonetic Spelling: ih·rith·ruh·blasts

- Mononuclear
Pronunciation link: https://www.merriam-webster.com/dictionary/mononuclear
IPA: /ˌmɑː.noʊˈnuː.kli.ər/, /ˌmɑː.noʊˈnjuː.kli.ər/
Phonetic Spelling: maa·noh·noo·klee·er / maa·noh·nyoo·klee·er

- Hematogones
Pronunciation link: https://www.merriam-webster.com/medical/hematogone
IPA: /ˈhiː.mət.əˌɡoʊnz/, /hɪˈmæt.əˌɡoʊnz/
Phonetic Spelling: hee·muht·uh·gohnz / hi·mat·uh·gohnz

- Isoclonic
Pronunciation link: No confirmed link found
IPA: /ˌaɪ.soʊˈklɑː.nɪk/
Phonetic Spelling: eye·soh·klaa·nik

- Immunophenotypes
Pronunciation link: https://www.merriam-webster.com/medical/immunophenotype
IPA: /ˌɪm.jə.noʊˈfiː.nəˌtaɪps/
Phonetic Spelling: im·yuh·noh·fee·nuh·types

- Cartography
Pronunciation link: https://www.merriam-webster.com/dictionary/cartography
IPA: /kɑrˈtɑːɡrəfi/
Phonetic Spelling: kar·taa·gruh·fee

- Leukemic
Pronunciation link: https://www.merriam-webster.com/dictionary/leukemia
IPA: /luːˈkiːmɪk/
Phonetic Spelling: loo·kee·mik

- Quantification
Pronunciation link: https://www.merriam-webster.com/dictionary/quantification
IPA: /ˌkwɑːn.tə.fəˈkeɪ.ʃən/
Phonetic Spelling: kwaan·tuh·fuh·kay·shuhn

- Physiological
Pronunciation link: https://www.merriam-webster.com/dictionary/physiological
IPA: /ˌfɪziəˈlɑːdʒɪkəl/
Phonetic Spelling: fiz·ee·uh·laa·ji·kuhl

- Ammonium
Pronunciation link: https://www.merriam-webster.com/dictionary/ammonium
IPA: /əˈmoʊniəm/
Phonetic Spelling: uh·moh·nee·uhm

- Giemsa
Pronunciation link: https://www.merriam-webster.com/dictionary/Giemsa
IPA: /ˈɡiːmzə/
Phonetic Spelling: geem·zuh

- Fluorescence
Pronunciation link: https://www.merriam-webster.com/dictionary/fluorescence
IPA: /flʊˈrɛsəns/
Phonetic Spelling: floo·res·uhns

- Immunostaining
Pronunciation link: No confirmed link found
IPA: /ˌɪm.jə.noʊˈsteɪ.nɪŋ/
Phonetic Spelling: im·yuh·noh·stay·ning

- Erythroblasts
Pronunciation link: https://www.merriam-webster.com/dictionary/erythroblast
IPA: /ɪˈrɪθ.rəˌblæsts/
Phonetic Spelling: ih·rith·ruh·blasts