# Submission ID #: 71304

# Scriptwriter Name: Poornima G

# Project Page Link: https://review.jove.com/account/file-uploader?src=21747768

# Title: Multicenter Flow Cytometry for Acute Myeloid Leukemia Measurable Residual Disease with LAIP/DfN and CD34+CD38− Leukemic Stem Cell-enriched Populations

# Authors and Affiliations:

# Adriana Plesa1,35*, Florent Dumezy2, Stéphanie Mathis3, Isabelle Arnoux4, Veronique Saada5  ,Valerie Bardet6 , Iole Oster1 ,Bouchra Badaoui7, Elsa Bera8, Edouard Bonneville9 , Boutheina Bouslama10 ,  Nicolas Chapuis11 ,Clementine Chauvel3 , Edouard Cornet12 , Louis Thomas Dannus13 , Camille Debord14 , Fréderic Feger15 , Jean Feuillard16  ,Thomas Fournet17, Francine Garnache17 , Franck Geneviève18, Delphine Gerard19, Valerie Goncalves Monteiro20, Julien Guy21 ,Veronique Harrivel22, Claire Hemar23, Cassandra Jadeau16, Caroline Mayeur-Rousse14,  Hélène Lapillonne10,  Veronique Latger-Cannard19 , Magali Le Garff-Tavernier24,  Pierre Lemaire3, Remi Letestu25  , Anne Catherine Lhoumeau26, Delphine Manzoni1, Aguirre Mimoun27 ,Vanessa Nivaggioni4,  Jennifer Osman28,  Victoria Raggueneau28 , Hubert Rambaud18, Tatiana Raskovalova9 , Anne Roggy17 , Mikael Roussel29 , Didier Sciortino26 , Valerie Soenen2, François Vergez30 ,Richard Veyrat Masson13, Jean Philippe Vial27 , Manon Vogrig31 , Orianne Wagner-Ballon7,37 , Jaja Zhu6 , Karine Celli-Lebras32, Raphaël Itzykson33,38, Christian Récher34, Hervé Dombret32,33,Claude Preudhomme2;36 , Christophe Roumier2,36*

# * Coordinators of the Acute Leukemia French Intergroup AML MRDflow & LSC network

# ¹Laboratory of Hematology and Flow cytometry, Lyon-Sud Hospital, Hospices Civils de Lyon, CHU-Lyon

# ²Laboratory of Hematology and Flow cytometry, CHU-Lille

# ³Laboratory of Hematology and Flow cytometry, Saint Louis Hospital, APHP

# ⁴Laboratory of Hematology and Flow cytometry, APHM-Marseille

# ⁵Laboratory of Hematology and Flow cytometry, IGR

# ⁶Laboratory of Hematology and Flow cytometry, Ambroise Pare University Hospital, APHP, Paris Saclay

# ⁷Département d’Hématologie et Immunologie Biologiques, Hôpitaux Universitaires Henri Mondor, APHP

# ⁸Laboratory of Hematology and Flow cytometry, CHU-Rouen

# ⁹Laboratory of Immunology and Flow cytometry, CHU-Grenoble

# ¹⁰Laboratory of Hematology and Flow cytometry, Trousseau Hospital, APHP

# ¹¹Laboratory of Hematology and Flow cytometry, Cochin Hospital, APHP

# ¹²Laboratory of Hematology and Flow cytometry, CHU-Caen

# ¹³Laboratory of Hematology and Flow cytometry, CHU-Clermont Ferrand

# ¹⁴Laboratory of Hematology and Flow cytometry, CHU-Nantes

# ¹⁵Laboratory of Hematology and Flow cytometry, Saint Antoine Hospital, APHP

# ¹⁶Laboratory of Hematology and Flow cytometry, CHU-Limoges

# ¹⁷Laboratory of Hematology and Flow cytometry, CHU-Besancon

# ¹⁸Laboratory of Hematology and Flow cytometry, CHU-Angers

# ¹⁹Laboratory of Hematology and Flow cytometry, CHU-Nancy

# ²⁰Laboratory of Hematology and Flow cytometry, CHU-Reims

# ²¹Laboratory of Hematology and Flow cytometry, CHU-Dijon

# ²²Laboratory of Hematology and Flow cytometry, CHU-Amiens

# ²³Laboratory of Hematology and Flow cytometry, CH-Valenciennes

# ²⁴Laboratory of Hematology and Flow cytometry, Pitie Salpetriere Hospital, APHP

# ²⁵Laboratory of Hematology and Flow cytometry, Avicenne Hospital, APHP

# ²⁶Laboratory of Hematology and Flow cytometry, IPC-Marseille

# ²⁷Laboratory of Hematology and Flow cytometry, CHU-Bordeaux

# ²⁸Laboratory of Hematology and Flow cytometry, Versailles Hospital

# ²⁹Laboratory of Hematology and Flow cytometry, CHU-Rennes

# ³⁰Laboratory of Hematology and Flow cytometry, CHU-Toulouse

# ³¹Laboratory of Hematology and Flow cytometry, CHU-Saint Etienne

# ³²Acute Leukemia French Association Group Coordination, IRSL

# ³³Department of Hematology, Saint Louis Hospital, APHP

# ³⁴Department of Hematology, CHU Toulouse, Université de Toulouse, Institut Universitaire du Cancer de Toulouse Oncopole

# ³⁵CRCL INSERM 1052/CNRS 5286, University of Lyon

# ³⁶UMR9020 CNRS-UMR-S1277 INSERM, University of Lille

# ³⁷Université Paris-Est Créteil (UPEC), INSERM U955 IMRB

# ³⁸Service Hématologie Adultes, Hôpital Saint Louis, APHP Paris, Université Paris Cité, INSERM, CNRS, Institut de Recherche Saint Louis (IRSL)

☐   All author names and affiliations are correct (city/state/country information not included in video title page).

# Corresponding Author:

# Adriana Plesa		adriana.plesa@chu-lyon.fr

# Email Addresses for All Authors:

Florent Dumezy 	 florent.dumezy@chu-lille.fr

Stéphanie Mathis 	stephanie.mathis@aphp.fr

Isabelle Arnoux 	 isabelle.arnoux@mail.ap-hm.fr

Veronique Saada 	veronique.saada@igr.fr

Valerie Bardet 	 valerie.bardet@aphp.fr

Iole Oster 	 	iole.oster@chu-lyon.fr

Bouchra Badaoui 	 bouchra.badaoui@aphp.fr

Elsa Bera 	 	Elsa.Bera@chu-rouen.fr

Edouard Bonneville 	 ebonneville@chu-grenoble.fr

Boutheina Bouslama 	boutheina.bouslama@aphp.fr

Nicolas Chapuis 	nicolas.chapuis@aphp.fr

Clementine Chauvel 	 clementine.chauvel@aphp.fr

Edouard Cornet 	 cornet-e@chu-caen.fr

Louis Thomas Dannus  ltdannus@chu-clermontferrand.fr

Camille Debord 	 camille.debord@chu-nantes.fr

Fréderic Feger 	frederic.feger@aphp.fr

Jean Feuillard 		jean.feuillard@unilim.fr

Thomas Fournet 	 tfournet@chu-besancon.fr

Francine Garnache 	 fgarnache@chu-besancon.fr

Franck Geneviève 	 FrGenevieve@chu-angers.fr

Delphine Gerard 	 D.GERARD@chru-nancy.fr

Valerie Goncalves Monteiro  vgoncalves-monteiro@chu-reims.fr

Julien Guy 		julien.guy@chu-dijon.fr

Veronique Harrivel 	 Harrivel.Veronique@chu-amiens.fr

Claire Hemar		 hemar-c@ch-valenciennes.fr

Cassandra Jadeau 	Cassandra.JADEAU@chu-limoges.fr

Caroline Mayeur-Rousse	 caroline.mayeurrousse@chu-nantes.fr

Hélène Lapillonne 	helene.lapillonne@aphp.fr

Veronique Latger-Cannard 	v.cannard@chru-nancy.fr

Magali Le Garff-Tavernier 	 magali.legarff@aphp.fr

Pierre Lemaire 	pierre.lemaire@aphp.fr

Remi Letestu 		remi.letestu@aphp.fr

Anne Catherine Lhoumeau 	LHOUMEAUAC@ipc.unicancer.fr

Delphine Manzoni 	delphine.manzoni@chu-lyon.fr

Aguirre Mimoun 	aguirre.mimoun@chu-bordeaux.fr

Vanessa Nivaggioni	 Vanessa.NIVAGGIONI@ap-hm.fr

Jennifer Osman 	josman@ch-versailles.fr

Victoria Raggueneau 	vraggueneau@ch-versailles.fr

Hubert Rambaud	 HuRambaud@chu-angers.fr

Tatiana Raskovalova 	TRaskovalova@chu-grenoble.fr

Anne Roggy		 aroggy@chu-besancon.fr

Mikael Roussel 	Mikael.ROUSSEL@chu-rennes.fr

Didier Sciortino 	 SCIORTINOD@ipc.unicancer.fr

Valerie Soenen 	Valerie.SOENEN@ chu-lille.fr

François Vergez 	Vergez.Francois@iuct-oncopole.fr

Richard Veyrat Masson 	rveyrat@chu-clermontferrand.fr

Jean Philippe Vial 	jean-philippe.vial@chu-bordeaux.fr

Manon Vogrig		 Manon.VOGRIG@chu-st-etienne.fr

Orianne Wagner-Ballon 	 orianne.wagnerballon@aphp.fr

Jaja Zhu 		jaja.zhu@aphp.fr

Karine Celli-Lebras 	karine.cellilebras@alfa-leukemia.org

Raphaël Itzykson 	raphael.itzykson@aphp.fr

Christian Recher	 Recher.Christian@iuct-oncopole.fr

Hervé Dombret	 herve.dombret@mac.com

Claude Preudhomme	Claude.PREUDHOMME@CHU-LILLE.FR

Christophe Roumier	Christophe.ROUMIER@chu-lille.fr

Adriana Plesa		adriana.plesa@chu-lyon.fr

## Author Questionnaire

1. Microscopy: Does your protocol require the use of a dissecting or stereomicroscope for performing a complex dissection, microinjection technique, or something similar?  Enter Yes or No.

If Yes, can you record movies/images using your own microscope camera?

Enter Yes or No.

If your protocol involves microscopy but you are not able to record movies/images with your microscope camera, JoVE will need to use our scope kit.

If your microscope does not have a camera port, the scope kit will be attached to one of the eyepieces and you will have to perform the procedure using one eye.

Enter make and model of microscope.

If a dissection or stereo microscope is required for your protocol, please list all shots from the script that will be visualized using the microscope (shots are indicated with the 3-digit numbers, like 2.1.1, 2.1.2, etc.).

Click here to list microscope shots, using the shot numbers from the protocol section of the video script.

2. Software: Does the part of your protocol being filmed include step-by-step descriptions of software usage?  Enter Yes or No.

If Yes, we will need you to record using screen recording software.

We recommend using the screen capture program OBS. JoVE’s tutorial for using OBS Studio is provided at this link: https://review.jove.com/v/5848/screen-capture-instructions-for-authors?status=a7854k

As these files are necessary for finalizing your script, please upload all screen-captured video files to your project page as soon as possible.

3. Filming location: Will the filming need to take place in multiple locations?   Enter Yes or No.

If Yes, how far apart are the locations? Click to enter distance between locations.

4. Testimonials (optional): Would you be open to filming two short testimonial statements live during your JoVE shoot? These will not appear in your JoVE video but may be used in JoVE’s promotional materials. Enter Yes or No.

If Yes, please provide the full name and position (e.g., Director of [Institute Name], Senior Researcher [University Name], etc.) of the author willing to participate.

Enter full author name, Enter author position

To ensure that your script can be filmed in one day, the protocol sections are cumulatively restricted to 55 shots (shots are the 3-digit numbers like 2.1.1, 2.1.2…etc)

Current Protocol Length

Number of Steps:  25

Number of Shots:  55 (35 SC)

# Interviews

Videographer: Obtain headshots for all authors available at the filming location.

Answers to these questions will become interview statements that you will deliver on camera.

- Introduction: Answer the 1st REQUIRED prompt and 1 additional question; Conclusion: Answer any 3 questions. No more than 5 interview statements will be included in the video.

- Enter the full name of the author who will deliver the statement.

- Speak naturally and avoid reading the lines.

- Limit the length of each statement to 20 words or fewer.

- Answers will be edited for length, clarity, and consistency with journal style guidelines.

Authors: Please note that the questions will not appear on screen. Please answer in stand-alone full sentence with sufficient context

INTRODUCTION:

REQUIRED: Summarize your research focus and the main questions it aims to answer.

  - Enter author name.: Click here to answer question. Please write in a style that you will be comfortable memorizing and speaking aloud. Limit length to 20 or fewer words.

Describe the research gap or limitation in existing methods and how this protocol addresses them.

  - Enter author name: Click here to answer question. Please write in a style that you will be comfortable memorizing and speaking aloud. Limit length to 20 or fewer words.

Describe the research contexts or systems in which this protocol can be applied.

  - Enter author name: Click here to answer question. Please write in a style that you will be comfortable memorizing and speaking aloud. Limit length to 20 or fewer words.

CONCLUSION:

What does this protocol allow researchers to study or measure?

  - Enter author name: Click here to answer question. Please write in a style that you will be comfortable memorizing and speaking aloud. Limit length to 20 or fewer words.

Describe the most important consideration or challenge when performing this protocol.

  - Enter author name: Click here if you choose this question. Please write in a style that you will be comfortable memorizing and speaking aloud. Limit length to 20 or fewer words.

Are there any additional methods or analyses that can be performed following this procedure?

  - Enter author name: Click here if you choose this question. Please write in a style that you will be comfortable memorizing and speaking aloud. Limit length to 20 or fewer words.

How can future studies build on this work? Describe the questions or research directions that can be investigated next.

  - Enter author name: Click here if you choose this question. Please write in a style that you will be comfortable memorizing and speaking aloud. Limit length to 20 or fewer words.

Videographer: Obtain headshots for all authors available at the filming location.

# Ethics Title Card

This research has been approved by the Human Ethics Committee of CHU Lyon and the Ethics Committee of the Institutional Review Board of Lille CHU

# Protocol

Please review this section to make sure that it accurately describes your protocol. Use Track Changes when making edits or revisions.

- The two-digit steps (e.g., 2.1., 2.2.) with purple font are the narration.  JoVE is responsible for the narration of the protocol and results.

- Red italics are pronunciation guides indicating how the word will be spoken.

- Filming should take no more than 10 minutes per step. If a step takes more than 10 minutes, prepare the product for that step in advance.

- The three-digit shots (e.g., 2.1.1., 2.2.2.) are the actions that the videographer will capture.

Videographer: Please always record an audio slate while filming each shot mentioned in the shotlist.

- Cytometer Standardization and Sample Preparation

Demonstrator: Click here to enter name of demonstrator(s)

  - To begin, prepare the 8-peak rainbow beads to determine cytometer-specific target values [1].

    - WIDE: Talent preparing the 8-peak rainbow beads. Videographer: Please obtain CU shot as well for this shot.

  - Set each machine according to the target values [1] and ensure that each channel's measured value is within 5% of its target [2]. Verify this setting systematically to maintain the longitudinal sensitivity of each platform [3].

    - Talent setting the machine according to the target values.

    - LAB MEDIA: Table 1.

    - Talent verifying the cytometer settings.

  - Collect 1 to 2 milliliters of bone marrow in EDTA at diagnosis and at monitoring time points, including minimal residual disease or MRD-1 through MRD4 and relapse [1]. Add 15 milliliters of ammonium chloride lysing solution to the bone marrow sample [2]. Mix gently and incubate for 10 minutes at room temperature [3].

    - Talent collecting an ethylenediaminetetraacetic acid bone marrow sample.

    - Talent adding ammonium chloride lysing solution to the bone marrow sample.

    - Talent gently mixing the sample and keeping it aside.

  - Then, centrifuge the sample at 450 g for 5 minutes at room temperature [1] and discard the supernatant [2]. Resuspend the cell pellet in 10 milliliters of PBS with 0.1% BSA [3]. Centrifuge again at 450 g for 5 minutes at room temperature [4] and discard the supernatant [5]. Resuspend the cell pellet at a concentration of 30 × 10⁶ white blood cells per milliliter in PBS [6].

    - Talent placing the sample in the centrifuge.

    - Talent discarding the supernatant.

    - Talent resuspending the cell pellet in PBS with 0.1% bovine serum albumin.

    - Talent placing the sample in the centrifuge.

    - Talent discarding the supernatant.

    - Talent resuspending the cell pellet in PBS.

- Immunostaining of Bone Marrow Cells

Demonstrator: Click here to enter name of demonstrator(s)

  - Distribute the monoclonal antibodies in tubes T1, T2, and T3 [1] according to the platform. Prepare an antibody table with the clone, volume, catalog reference, and supplier for each antibody [2].

    - Talent distributing monoclonal antibodies into tubes T1, T2, and T3.

    - LAB MEDIA: Table 2.

  - Add 2 to 3 × 10⁶ lysed cells per tube in 100 microliters of cell suspension [1] and incubate the cells with antibodies for 15 minutes in the dark at room temperature [2].

    - Talent adding lysed cells to each tube.

    - Talent placing the tubes in the dark for incubation at room temperature.

  - Wash the cells with 3 milliliters of PBS [1]. Centrifuge at 450 g for 5 minutes [2] and discard the supernatant [3]. Resuspend the stained cells in 300 microliters of PBS [4].

    - Talent washing the cells with PBS.

    - Talent placing the tubes in the centrifuge.

    - Talent discarding the supernatant.

    - Talent resuspending the stained cells in PBS.

- Flow Cytometry Acquisition and LAIP Analysis

Demonstrator: Click here to enter name of demonstrator(s)

  - Open the MRD assay or acquisition protocol with the adapted settings and compensation matrix for MRD leukemia-associated immunophenotype, different from normal, and leukemic stem cell measurement [1].

    - SCREEN: MRD assay or acquisition protocol being opened with the adapted settings and compensation matrix.

Authors: Please create screen capture videos of the shots labeled as SCREEN, write a screenshot summary, and upload the files to your project page as soon as possible: https://review.jove.com/account/file-uploader?src=21747768

  - Load the samples [1] and acquire a minimum of 500,000 to 1,000,000 white blood cells in each tube at diagnosis and in the MRD follow-up samples to improve clustering and confidence when detecting low-level MRD [2].

    - Talent loading the samples.

    - SCREEN: Acquisition in process.

  - Use a common template and dot-plot sequence that incorporates a predefined harmonized gating strategy adapted to the analysis software used at each center [1].

    - SCREEN: Show the common template and dot-plot sequence with the predefined harmonized gating strategy.

  - Capture leukemia-associated immunophenotype or LAIP patterns in bulk blasts by drawing a single-cell gate based on forward scatter width versus forward scatter height [1].

    - SCREEN: the single-cell gate being drawn on forward scatter width versus forward scatter height.

  - Draw a live-cell gate based on forward scatter versus side scatter [1].

    - SCREEN: live-cell gate being drawn on forward scatter versus side scatter.

  - Then, draw a blast gate on CD45 dim and side scatter [1], followed by the primitive marker gate using CD34 and CD117 [2].

    - SCREEN: blast gate being drawn on CD45 dim and side scatter.

    - SCREEN: primitive marker gate being drawn using CD34 and CD117.

  - Using fixed biparametric dot plots, draw the relevant gates, including LAIP 1 through LAIP 8 [1]. After gating the physiological population, use the Boolean operators for the LAIP approach [2-TXT].

    - SCREEN: LAIP gates being drawn on fixed biparametric dot plots.

    - SCREEN:  Show the data after applying the LIAP approach. TXT: Refer to the text protocol for the LAIP approach Boolean operators

- Different-From-Normal (DfN) Analysis

Demonstrator: Click here to enter name of demonstrator(s)

  - Define abnormal populations when the cluster of abnormal cells is separated from normal cells by at least 0.5 decade on the log scale [1]. To capture a different-from-normal pattern in bulk blasts, draw live cells using forward scatter and side scatter [2].

    - SCREEN: the abnormal cell cluster being separated from the normal cells on the log-scale plot.

    - SCREEN: Show live cells being gated using forward scatter and side scatter.

  - Draw single cells by removing doublets based on the forward scatter peak and the side scatter area [1]. Then, draw the mononuclear cell gate using forward scatter and side scatter [2].

    - SCREEN: Show the single-cell gate being drawn after removing doublets based on the forward scatter peak and the side scatter area.

    - SCREEN: Show the mononuclear cell gate being drawn using forward scatter and side scatter.

  - Apply a cleanup gate to remove lymphocytes using CD45, CD36, and side scatter [1]. Next, apply a cleanup gate to remove erythroblasts and debris using CD45 and CD36 [2].

    - SCREEN: Show the cleanup gate being applied to remove lymphocytes using CD45, CD36, and side scatter.

    - SCREEN: Show the cleanup gate being applied to remove erythroblasts and debris using CD45 and CD36.

  - Use two input gates, CD34 positive, CD45, and side scatter, and CD117 positive, CD45, and side scatter, to define the empty boxes [1].

    - SCREEN: Show the two input gates being used to define the empty boxes.

  - Draw each different-from-normal empty box based on internal control populations, including lymphocytes for cross-lineage markers and normal myeloid precursors for loss, decrease, or overexpression markers [1]. Compare the empty boxes with reference bone marrow samples, including regenerating, inflammatory, and healthy donor bone marrow [2].

    - SCREEN: Show each different-from-normal empty box being drawn based on the internal control populations.

    - SCREEN: Show the empty boxes being compared with the reference bone marrow samples.

- Leukemic Stem Cell Analysis and CD38 Threshold Definition

Demonstrator: Click here to enter name of demonstrator(s)

  - Define the leukemic stem cell or LSC input gate as mononuclear cell CD34-positive using the gating sequence [1]. Define the CD34-positive CD38-negative fraction P6 using a rigorous threshold based on the CD38 fluorescence-minus-one level [2-TXT]. Use CD36 to clean up nonspecific events in the final MRD LSC gate [3]

    - SCREEN: leukemic stem cell input gate being defined as mononuclear cell CD34-positive using the gating sequence.

    - SCREEN: Show the P6 fraction being defined using the CD38 fluorescence-minus-one threshold. TXT: Refer to the text protocol for the Boolean operators

    - SCREEN: cleaning up non-specific events in final MRD LSC gate

  - Next, set the CD38 channel high enough to distinguish CD38-negative populations from CD38-positive populations with good resolution [1].

    - SCREEN: Show the CD38 channel being adjusted to separate the populations.

  - Set the CD38 channel based on CD38 bright populations [1]. Identify plasma cells as the brightest population and hematogones according to the measurement range of each platform [2].

    - SCREEN: Show the CD38 channel set based on CD38 bright populations.

    - SCREEN: Show plasma cells identified as the brightest population and hematogones identified according to the platform measurement range.

  - Then, define the CD34-positive CD38-negative population P6 using the CD38 fluorescence-minus-one approach validated by an isoclonic CD38 control [1]. Define P7, the CD34-positive and CD38-low population, using the hematogone-negative level [2]. Define P8, the CD34 positive and CD38 high population, using the hematogone positivity level and plasma-cell negative level [3].

    - SCREEN: Show the P6 population being defined using the validated CD38 fluorescence-minus-one approach.

    - SCREEN: Show P7 being defined using the hematogone negative level.

    - SCREEN: Show P8 being defined using the hematogone positivity level and plasma-cell negative level.

  - Use the same CD38-negative threshold across the harmonized group, with robust methods such as CD38 fluorescence minus one or an isoclonic CD38 control [1].

    - SCREEN: Show the same CD38 negative threshold being applied across the harmonized group.

  - Use at least one aberrant immunophenotype in the CD34-positive CD38-negative fraction to define LSC-enrichment [1]. Include LSC markers or lineage markers [2-TXT].

    - SCREEN: Show an aberrant immunophenotype being used in the CD34-positive CD38-negative fraction.

    - SCREEN: Show the leukemic stem cell and lineage markers included for analysis. TXT: LSC markers: CLL1/TIM3/CD97/CD45RA/CD123

Lineage markers: CD7/CD19/CD56/CD33

  - Apply the assay sensitivity thresholds, with a limit of quantification of 10⁻⁴ and a limit of detection of 10⁻⁵ [1]. Define LSC positivity at diagnosis as at least 1% of CD45 and side scatter blasts, and in MRD follow-up as at least 0.01% of white blood cells [2]. Finally, provide the harmonized clinical flow report [3].

    - SCREEN: Show the assay sensitivity thresholds being applied.

    - SCREEN: Show the leukemic stem cell positivity thresholds being applied.

# Results

Please review this section to make sure that it accurately reflects your findings.

- This section will not be recorded by the videographer. It only includes the figures/tables from your manuscript (called LAB MEDIA).

- Use Track Changes when making edits or revisions. Ensure the voiceover length is below 200 words. Current word count: XXX.

- Please note that the video cannot include voiceover without an accompanying visual.

- Results

  - After applying the defined settings, the cell staining patterns overlapped across the 4 flow cytometer platforms [1], with the main difference being a translation of the cell clouds along the logarithmic scales [2].

    - LAB MEDIA: Figure 6. Video editor: Show the 4 platform panels side by side and highlight the matching cell pattern shapes across the plots.

    - LAB MEDIA: Figure 6. Video editor: Emphasize that the cell clouds appear shifted left or right between the platforms while keeping the same overall shape.

  - The wet external quality assessment showed reproducible distributions of %P6 CD34-positive CD38-negative cells in white blood cell CD45+ cells across the 5 assessment rounds [1].

    - LAB MEDIA: Figure 7A. Video editor: Highlight the violin plots representing the 5 assessment rounds for %P6 CD34+CD38− cells.

  - The wet external quality assessment showed reproducible distributions of % mononuclear cell CD34-positive cells in white blood cell CD45 positive cells across the 5 assessment rounds [1].

    - LAB MEDIA: Figure 7B. Video editor: Highlight the violin plots representing the 5 assessment rounds for % mononuclear cell CD34+ cells.

  - More than 90% of the locally analyzed data were included within the centralized interval based on the series mean in DryEQA2017 [1] and DryEQA2024 [2].

    - LAB MEDIA: Figure 8A. Video editor: Highlight the participant results clustered around the central reference line in the DryEQA2017 graph.

    - LAB MEDIA: Figure 8B. Video editor: Highlight the participant results clustered around the central reference line in the DryEQA2024 graph.

  - The leukemia-associated immunophenotype cartography established at diagnosis [1] was applied to the follow-up sample to quantify measurable residual disease [2].

    - LAB MEDIA: Figure 9A. Video editor: Highlight the black cell population marked as the leukemia-associated immunophenotype in the diagnosis panels.

    - LAB MEDIA: Figure 9A. Video editor: Highlight the black measurable residual disease population in the follow-up panels.

  - The different-from-normal approach identified abnormal cells in the areas not observed in normal bone marrow [1].

    - LAB MEDIA: Figure 9B. Video editor: Highlight the boxed regions that are empty in the normal plots and contain black abnormal cells in the follow-up plots.

  - The leukemic stem cell-enriched cells expressing the defined marker profile were identified at diagnosis within the CD34-positive CD38-negative fraction [1], and the measurable residual disease leukemic stem cell-enriched cells with the same profile persisted at follow-up [2].

    - LAB MEDIA: Figure 10A. Video editor: Highlight the orange cell population in the diagnosis sample.

    - LAB MEDIA: Figure 10B. Video editor: Highlight the orange cell population in the follow-up sample.

  - The limit of blank for leukemic stem cell detection was less than 0.0005% [1], corresponding to 14 events in the leukemic stem cell Boolean gate [2].

    - LAB MEDIA: Figure 11A. Video editor: Highlight the value showing a limit of blank below 0.0005%.

    - LAB MEDIA: Figure 11A. Video editor: Highlight the boxed cluster labelled with 14 events.

  - The serial dilution of leukemic stem cell-enriched cells showed a correlation between the measured measurable residual disease values and the theoretical values [1], corresponding to a limit of quantification of 0.002% [2].

    - LAB MEDIA: Figure 11B. Video editor: Highlight the points aligned along the correlation line in the graph.

    - LAB MEDIA: Figure 11B. Video editor: Highlight the value indicating a limit of quantification of 0.002%.

  - The serial dilution of blasts with different leukemia-associated immunophenotypes showed a correlation between the measured measurable residual disease values and the theoretical values [1], with a limit of quantification below 0.1% [2].

    - LAB MEDIA: Figure 12A. Video editor: Highlight the data points following the trend line in the dilution graphs.

    - LAB MEDIA: Figure 12A. Video editor: Highlight the area indicating a limit of quantification below 0.1%.

  - The limit of quantification reached the 0.1% clinical threshold for all tested leukemia-associated immunophenotypes [1].

    - LAB MEDIA: Figure 12B. Video editor: Highlight the blue bars and the red horizontal line marking the 0.1% clinical threshold.

Watch this video and apply the settings while capturing the screen to obtain high-quality videos: https://www.youtube.com/watch?v=Y67dM1kKcTE

NOTE to the Authors:

You have two options for filming and submitting the SCREEN footage:

Option 1: Uploading Separate Clips (Highly Recommended)

Please upload a separate video file for each SCREEN shot (tagged as SCREEN), naming the file according to the shot number (e.g., 2.4.3, 1.5.1, 3.2.1, 3.1.3, etc.).

- Each clip must be a maximum of 20 to 25 seconds to match the voice narration.

- If an action takes longer, provide only the best, most representative 20-25 second segment.

- Please do not provide longer clips as we won’t be able to use them.

Option 2: Uploading a Single Continuous Video

If you prefer to record all steps together and upload a single video file, that is acceptable. In this case, you must write the precise timestamp for each shot within the script.

For example, if the action described in shot 2.13.1 is present in the file xxx.mp4 and occurs between timestamps 00:30 and 00:45 , then write the following in the script after each shot description or as a comment bubble.

(write the file name and timestamp after each shot): [File name] [Start Time]-[End Time]

- Example 	2.13.1  SCREEN: performing ...........action. xxx.mp4 00:30-00:45

2.13.2, SCREEN: cutting the…..........action. xxx.mp4 01:10-01:20

- The duration between the start and end times for each shot must be limited to 20 to 25 seconds.

- If a shot is longer than 25 seconds, choose two short segments (one at the beginning and one at the end) that collectively do not exceed 25 seconds.

- Example: 3.1.1 SCREEN: inserting a needle... xxx.mp4 00:00-00:10 and 02:30-02:40