- Time-Lapse Imaging By Automated Live-cell Imaging System

Demonstrator: Rong-Chen Chang

  - Switch on the imaging system, environmental control module, and carbon dioxide gas supply at least 1 hour before imaging [1]. Open the Environment Control panel in the automated live-cell imaging software [2]. Set the chamber temperature to 37 degrees Celsius and carbon dioxide to 5 percent ensuring that both settings are stabilized before loading the plate [3].

    - Talent switching on the imaging system and environmental control module.

    - SCREEN: Show the automated live-cell imaging software interface and opening the Environment Control panel. 00:00 - 00:01

SCREEN: Show the user entering 37 degrees Celsius temperature and 5% CO2 settings and show the stabilized settings. *shown in 4.6.1 (01:25 - 01:30)

  - Place the 24-well plate into the imaging chamber of the imaging system [1].

    - Talent loading the 24-well plate into the imaging chamber.

  - For initial configuration, open the Task Manager in the automated live-cell imaging software [1]. Select Imager manual mode [2]. Click Capture now to initialize imaging parameters [3]. Select a 10X objective [4].

    - SCREEN: Show the Task Manager window in the automated live-cell imaging software. 00:00 - 00:01

    - SCREEN: Show the user selecting Imager manual mode. 00:02- 00:03

    - SCREEN: Show the user clicking Capture now. 00:04 - 00:011

    - SCREEN: Show the 10× objective selected in the software. 00:012 – 00:13

  - Choose microplate as the filter setting [1]. Select 24-well plate from the plate library of the automated live-cell imaging software [2]. For focus and image setup, select the Phase Contrast channel [3]. Select a representative well such as A1 [4]. Activate autofocus before time-lapse acquisition [5].

    - SCREEN: Show microplate selected as the filter setting. 00:14 - 00:17

    - SCREEN: Show the user selecting a 24-well plate from the plate library.

00:18 - 00:39

SCREEN: Show the Phase Contrast channel selected. 00:40 - 00:43

SCREEN: Show representative well A1 selected on the plate map. 00:44 - 00:45

    - SCREEN: Show autofocus being activated before acquisition. 00:46 - 00:55

  - Capture an image using the camera icon [1]. Confirm cell morphology and focus quality [2]. Navigate to Process/Analyze [3]. Select Create experiment from an image set to configure the experiment [4].

    - SCREEN: Show the user capturing an image using the camera icon. 00:56 - 01:06

    - SCREEN: Show the captured image with proper cell morphology and focus quality. 01:07 - 01:10

    - SCREEN: Show the user navigating to Process/Analyze. 01:11 - 01:17

    - SCREEN: Show the user selecting Create experiment from an image set.

01:18 - 01:22

  - To set the time-lapse acquisit ion parameters, open the Procedure settings panel [1]. Set the temperature to 37 degrees Celsius and carbon dioxide to 5 percent [2]. Under Start Kinetic, set the run time to sixteen hours [3].

SCREEN: Show the Procedure settings panel. 01:23 - 01:24

    - SCREEN: Show the temperature and carbon dioxide settings entered in the panel.

01:25 - 01:30

SCREEN: Show the user entering a run time of 16 hours in the Start Kinetic settings. 01:31 - 01:34

  - Set the imaging interval to ten minutes [1]. Set the total reads to 97 [2]. Use an imaging interval of 10 to 15 minutes to balance temporal resolution and phototoxicity while capturing early phagocytic events [3].

    - SCREEN: Show the imaging interval set to 10 minutes. 01:35 - 01:38

    - SCREEN: Show the total reads value set to 97. 01:39 - 01:40

    - SCREEN: Show the finalized experiment setup with all kinetic imaging parameters displayed. 01:40 - 01:45