71542

71542_screenshot_1

    - Set senter field of 335 millitesla, sweep width of 10 millitesla, is being set in the spectrometer software. 0:07-0:20

71542_screenshot_2

    - A modulation amplitude of 500 microtesla and a microwave power of 1.6 milliwatts (18dB) is being entered in the spectrometer software. 0:02-0:15

71542_screenshot_3

    - The number of scans is being set to 32, sweep time is being set to 11 seconds and time constant is being set to 0 in the spectrometer software. 0:02-0:12

71542_screenshot_4

    - Click Set parameters to transfer parameters to spectrometer, then click Tune 0:02-0:12

71542_screenshot_5

    - Start the acquisition ( Start Scan) and display the electron paramagnetic resonance spectrum being recorded. 0:01-0:40

MB-Screen-4.1.1

Open the previously saved spectrum that will be used for multiharmonic analysis.

MB-Screen-4.1.2

Enable the dual-channel acquisition mode in the software.

MB-Screen-4.2.1

Depending on the number of acquired harmonic frequencies, limit the number of harmonics included in the analysis. In this example, set the value to 30.

MB-Screen-4.2.2

Adjust the RF phase to obtain the optimal phase setting.

MB-Screen-4.3.1

Adjust the modulation phase and display the pure absorption signal in Channel 1.

MB-Screen-4.3.2

Adjust the second modulation phase and display the pure dispersion signal in Channel 2.

MB-Screen-4.4.1

Adjust the amplitude of the second modulation until the narrowest spectral linewidth is obtained without introducing additional distortions.

MB-Screen-4.5.1

Adjust the low-pass filter settings so that the signal shape remains undistorted while minimizing digital noise.

MB-Screen-4.6.1

Open the Noise Suppression tab and enable the Noise Suppression option.

MB-Screen-4.6.2

Optimize the Smoothing Factor and Narrow Line Refinement parameters. Their optimal values depend primarily on the spectral linewidth: the narrower the line, the greater the benefit of these corrections. For relatively broad lines, as in this example, disabling Narrow Line Refinement may provide better results.

MB-Screen-4.8.1

Save the processed data in the desired file format for further analysis.

MB-Screen-4.8.2

Import the processed data into analysis software (e.g., Origin) and perform the double integration procedure.