# Submission ID #: 71605

# Scriptwriter Name: Pallavi Sharma

# Project Page Link: https://review.jove.com/account/file-uploader?src=21839573

# Title: Methodological Implementation of Small Animal Irradiator Systems for Image-Guided Preclinical Studies

# Authors and Affiliations:

# Guilherme R. Ferreira*, Livia Huang*, Lori Strong, Geoffrey D. Hugo, Buck E. Rogers, Arash Darafsheh

# Department of Radiation Oncology, WashU Medicine

# *These authors contributed equally

# Corresponding Authors:

# Arash Darafsheh		(arash.darafsheh@wustl.edu)

# Email Addresses for All Authors:

# Arash Darafsheh		(arash.darafsheh@wustl.edu)

# Guilherme R. Ferreira 		(f.guilherme@wustl.edu)

# Livia Huang 			(livia@wustl.edu)

# Lori Strong 			(strongl@wustl.edu)

# Geoffrey D. Hugo 		(gdhugo@wustl.edu)

# Buck E. Rogers 		(b.rogers@wustl.edu)

## Author Questionnaire

1. Microscopy: Does your protocol require the use of a dissecting or stereomicroscope for performing a complex dissection, microinjection technique, or something similar?  No

2. Software: Does the part of your protocol being filmed include step-by-step descriptions of software usage?  Yes

3. Filming location: Will the filming need to take place in multiple locations?   Yes

Same building, different floors. Less than half a mile apart.

4. Testimonials (optional): Would you be open to filming two short testimonial statements live during your JoVE shoot? These will not appear in your JoVE video but may be used in JoVE’s promotional materials. No

Current Protocol Length

Number of Steps:  25

Number of Shots:  63

# Interviews

Videographer: Obtain headshots for all authors available at the filming location.

Authors: Please note that the questions will not appear on screen. Please answer in stand-alone full sentence with sufficient context

INTRODUCTION:

REQUIRED: Summarize your research focus and the main questions it aims to answer.

  - Buck Rogers: Our research uses a small animal irradiator to deliver precise, clinically relevant radiation doses to rodents and investigate the resulting radiation effects.

    - INTERVIEW: Named talent says the statement above in an interview-style shot, looking slightly off-camera. Suggested B roll: 2.1.2

Describe the research contexts or systems in which this protocol can be applied.

  - Buck Rogers: This protocol can be used for cancer research, normal tissue radiation toxicity studies, radiation biology, as well as other radiation effects.

    - INTERVIEW: Named talent says the statement above in an interview-style shot, looking slightly off-camera.

CONCLUSION:

What does this protocol allow researchers to study or measure?

  - Guilherme Rosa Ferreira: The protocol shows researchers how to deliver radiation to small animals in the most precise way possible.

    - INTERVIEW: Named talent says the statement above in an interview-style shot, looking slightly off-camera.

Describe the most important consideration or challenge when performing this protocol.

  - Guilherme Rosa Ferreira: The key considerations while performing this protocol are verifying the correct filter and properly positioning the isoflurane tube in the animal’s mouth to ensure reliable experiments.

    - INTERVIEW: Named talent says the statement above in an interview-style shot, looking slightly off-camera.

Are there any additional methods or analyses that can be performed following this procedure?

  - Buck Rogers: Following rodent irradiation, additional analyses can include monitoring rodents for radiation effects and harvesting tissue samples for subsequent evaluation.

    - INTERVIEW: Named talent says the statement above in an interview-style shot, looking slightly off-camera.

Videographer: Obtain headshots for all authors available at the filming location.

# Ethics Title Card

This research has been approved by the Institutional Animal Care and Use Committee (IACUC) at Washington University in St. Louis.

# Protocol

Videographer: Please always record an audio slate while filming each shot mentioned in the shotlist.

- Performing Cone-Beam CT Imaging for Mice with Radiator System 2

Demonstrator: Guilherme Rosa Ferreira

  - After initializing and warming up the radiator system, open the cabinet door [1] and insert the 2-millimeter aluminum filter into the filter slot if it is not already in place [2]. Secure the couch to the platform by depressing the quick-release pin and aligning the pin with the hole in the center of the platform [3].

    - WIDE: Talent opening the cabinet door.

    - Talent inserting the 2 millimeter aluminum filter into the filter slot.

    - Talent depressing the quick-release pin and securing the couch to the platform.

  - Next, place the rodent bed onto the couch and secure it with tape [1]. Next, attach the isoflurane tube to the head of the rodent bed [2] and turn the gas on to administer 1 to 3% dose of anesthesia [3]. Place the anesthetized mouse onto the bed in a prone or supine position according to the experiment protocol [4]. NOTE: The VO is edited for the additional shot

    - Talent placing the rodent bed onto the couch and securing it with tape.

    - Talent attaching the isoflurane tube to the head of the bed.

      - Talent turning the gas on. Added shot

    - Talent placing the mouse onto the rodent bed in the required position.

  - Then, open Software 5. Under the Radiation delivery category, select CT targeting with either the manual or treatment protocol [1].

    - SCREEN:  71605_screenshot_1.mp4: 00:00-00:12

  - Under Preset selection, choose the preset according to the target tissue type and select Next [1]. Verify that the appropriate filter for the planned beam quality is inserted and confirm that the displayed filter status is green [2]. Select Start to begin the CT scan, and the shutter will automatically open and close [3].

    - SCREEN: 71605_screenshot_1.mp4: 00:13-00:17, 00:21-00:22, 00:26-00:28, 00:37-00:43

    - SCREEN: 71605_screenshot_1.mp4:   00:47-00:57

    - SCREEN:71605_screenshot_1.mp4:  00:58-01:02

    - SCREEN: Shot of the shutter closing and opening  NOTE: Not provided

  - Then, use the webcam to ensure that the gantry will not collide with the couch or animal as it rotates 360 degrees [1]. If a collision looks imminent, press the red emergency stop button on the console to halt the gantry [2]. Once the CT scan is complete, select Next [3]. Review the acquired CT scan in the coronal, sagittal, and axial panels, then zoom in and scroll through the acquired slices to adjust the views as necessary [4].

    - SCREEN: 71605_screenshot_1.mp4:  01:28-02:00

    - Talent pressing the red emergency stop button on the console.

    - SCREEN: 71605_screenshot_1.mp4:  02:47-02:50

    - SCREEN: 71605_screenshot_1.mp4:   02:50-03:00

  - Use the mouse to drag the crosshair to the target isocenter [1]. Select Move Stage. If necessary, select Scan Again to verify the target position and click Next [2]. On the Export page, select Go to export the DICOM (Die-Com) file [3].

    - SCREEN: 71605_screenshot_1.mp4:  03:25-03:30, 03:39-03:42

    - SCREEN: 71605_screenshot_1.mp4:  03:46-04:00

    - SCREEN:  71605_screenshot_1.mp4:  06:27-6:36

- Generation and Evaluation of the Radiation Treatment Plan for System 2

Demonstrator: Lori Strong

  - Open Software 6. Select the folder containing DICOM files and select Import Selected Data [1].

    - SCREEN: 71605_screenshot_2.mp4:   00:02-00:07, 00:23-00:25

  - Once the DICOM file is loaded, select the CT2MD (C-T-Two-M-D) tab and slide the density bar until the rodent tissue is visible [1].

    - SCREEN: 71605_screenshot_2.mp4:   00:29-00:33, 00:39-00:49

  - Next, select the Contouring tab. To add contours, select Add, name the contour, and specify its type under the Type drop-down menu [1].

    - SCREEN: 71605_screenshot_2.mp4:  00:50-01:00

  - Select the desired tool, then select Automatic and brush or outline the desired organ [1]. Use the scrollbar to move through the frames and continue outlining the organ to increase interpolation accuracy [2]. Once the last frame of the organ is outlined, select Interpolate [3].

    - SCREEN: 71605_screenshot_2.mp4:  01:01-01:10

    - SCREEN: 71605_screenshot_2.mp4:  01:29-01:38

    - SCREEN: 71605_screenshot_2.mp4:  01:39-01:41

  - Review the frames to confirm that the interpolation outlined the organ as expected and correct any frames that have deviated [1]. Select the Planning tab and select Add to add an isocenter to the plan. Then align the isocenter to the desired target in the coronal and axial planes and enter the expected dose [2].

    - SCREEN: 71605_screenshot_2.mp4:   01:45-01:54,

    - SCREEN: 71605_screenshot_2.mp4:   02:16-02:28, 02:41-02:54

  - To add beams, select Add under the Beams section. [1]. Choose between static and dynamic beams, select the collimator according to the planning parameters, adjust the beam angles as desired, and select a couch angle if necessary [2]. Under Calculation, select Start Calculations [3].

    - SCREEN: 71605_screenshot_2.mp4:  02:55-02:59

    - SCREEN: 71605_screenshot_2.mp4:  03:03-03:06, 03:17-03:18, 03:28-03:38,  03:43-03:53

    - SCREEN: 71605_screenshot_2.mp4:   04:05-04:10, 04:35-04:37

  - Once the calculation is complete, review the Volume Slice display [1]. Select the tabs to display the Dose Volume Histogram and Dose Volume Metrics [2].

    - SCREEN: 71605_screenshot_2.mp4:  04:39-04:42

    - SCREEN:  71605_screenshot_2.mp4:  04:45-04:56

  - Under Evaluation, verify the beam-on time, dose to the isocenter, Dose Volume Histogram, and Dose Volume Metrics [1]. Select Export data, name the file, and select Export to file to generate an initialization file [2].

    - SCREEN: 71605_screenshot_2.mp4:   05:05-05:06, 05:11-05:13, 05:17-05:18, 05:25-05:30

    - SCREEN: 71605_screenshot_2.mp4:   05:34-05:46.

- Radiation Treatment Using Radiator System 2

  - After opening the door, remove the 2-millimeter aluminum filter and insert the 0.3-millimeter copper filter and the collimator according to the experiment protocol [1]. Navigate back to the System 2 Program [2].

    - Talent removing the 2 millimeter aluminum filter and inserting the 0.3 millimeter copper filter and the collimator.

    - SCREEN: System 2 Program being opened.  NOTE: Not provided

  - For a treatment protocol, in the System 2 Program, go to the Protocol Selection tab, select Load Protocol, and choose the initialization file [1]. Verify the plan name and parameters by selecting the Show protocol details box. Then, select Next [3].

    - SCREEN: 71605_screenshot_3.mp4:    00:00-00:06

    - SCREEN: 71605_screenshot_3.mp4:    00:07-00:26

  - After that, select Start Treatment followed by Start Beam [1]. If more than one beam is included, select Start Beam again after the previous beam is delivered to begin the next beam [2]. Watch the webcam to ensure that there are no collisions [3-TXT].

    - SCREEN: 71605_screenshot_3.mp4:   00:30-00:40.

    - SCREEN: 71605_screenshot_3.mp4:   01:29-01:33,

    - SCREEN: 71605_screenshot_3.mp4:    01:45-01:51, 02:00-02:08 TXT: Refer to the text protocol for the post-treatment procedure

- Acquisition of Cone-Beam CT Images with Radiator System 1

Demonstrator: Livia Huang

  - After initializing and warming up the radiator system, open the cabinet door [1]. Slide the CT (C-T) window cover up and pull the CT unit forward until it locks [2]. Ensure that all couches are removed from the platform [3]. Remove the collimator and holder from the gantry head [4]. Then insert the 1-millimeter aluminum filter for imaging [3].

    - WIDE: Talent opening the System 1 cabinet door.

    - Talent sliding the CT window cover upward and pulling it forward until it locks.

    - Talent checking that all couches have been removed from the platform.  NOTE: 5.1.3 and 5.1.4 are not filmed

    - Talent removing the collimator and holder from the gantry head.

    - Talent inserting the 1 millimeter aluminum filter into the gantry head.

  - Once the detector calibration is completed, place the couch onto the platform and secure it by tightening the screw [2]. NOTE: The VO is edited for the deleted shot

    - WIDE: Talent opening the cabinet door. NOTE: Not filmed

    - Talent placing the couch onto the platform and tightening the screw.

  - Position the mouse onto the bed in a prone or supine position as demonstrated earlier [3].

    - Talent placing the mouse onto the rodent bed in the required position.

  - In Software 1, under Imaging and Robotics Control, select CBCT Acquisition (C-B-C-T Acquisition) for the principal investigator and study [1]. Name the scan, select the desired reconstruction protocol, set the number of 2-dimensional projections and the couch speed, then select OK [2]. Watch the webcam on the software screen to ensure there are no collisions. If a collision looks imminent, press Stop to halt the gantry [3].

    - SCREEN: 71605_screenshot_4.mp4:    00:16-00:33

    - SCREEN: 71605_screenshot_4.mp4:   : 00:32-00:38, 00:56-01:00, 01:04-01:15

    - SCREEN: 71605_screenshot_4.mp4:     01:54-02:00

- Execution of the Radiation Treatment Using System 1

  - Once the segmentation threshold is adjusted, select Treatment Planning. To create a new plan, under Targets, press the green plus sign for each desired isocenter and move the isocenter to the target [1-TXT].

    - SCREEN: Treatment three views.Planning being selected. Isocenters being moved to the center of the target organ and adjusted in all

    - SCREEN:  71605_screenshot_5.mp4:   00:00-00:11 TXT: Refer to the text protocol for the segmentation threshold workflow

  - For each target, set the dose by double-clicking the Dose field [1]. Under Beams, press the green plus sign with a B to add a static beam [2]. Set the beam weight, collimator, and gantry angle for each beam [3]. To add a beam for a different isocenter, select the desired isocenter under Targets so that it is highlighted in purple, and then add the beam. [4]. Under Dose, select the desired Dose Engine using Superposition Convolution or Monte Carlo, then select Compute Dose [5].

    - SCREEN: 71605_screenshot_5.mp4:   00:12-00:15.

    - SCREEN: 71605_screenshot_5.mp4:    00:16-00:18

    - SCREEN: 71605_screenshot_5.mp4:   00:19-00:38

    - SCREEN: 71605_screenshot_5.mp4:   00:42-00:52, 00:58-01:16

    - SCREEN: 71605_screenshot_5.mp4:   01:17-01:36

  - Select Verify, then click Compute Isodose and Compute DVH (Compute-Dose-Volume-Histogram) to ensure the plan and dose-volume histograms accurately deliver the desired dose to the targets and organs at risk [1].

    - SCREEN: 71605_screenshot_5.mp4: 01:38-01:41, 03:17-03:20,  04:09-04:20

  - After opening the cabinet door, remove the aluminum filter and insert the copper filter [1]. Pull the CT window cover down and push it back until it locks [2]. Insert the collimator into the holder and slide the unit back into the gantry head [3]. Select Execute and Execute Beam [4].

    - Talent removing the aluminum filter and inserting the copper filter.

6.4.3	   Talent pulling the computed tomography window cover down and pushing it back until it locks. NOTE: 6.4.3 and 6.4.2. are swapped

6.4.2	   Talent inserting the collimator into the holder and sliding the unit back into the gantry head.

6.4.4	   SCREEN: 71605_screenshot_5.mp4: 04:28-04:37

Video editor: please use title card to show the following information

- Remove the anesthetized mouse from the machine after treatment.

- Observe until fully awake and recovered from anesthesia.

- Monitor once daily for 3 days according to protocol.

NOTE: This script does not have results

- Radiator
Pronunciation link: https://www.merriam-webster.com/dictionary/radiator
IPA: /ˈreɪdiˌeɪtər/
Phonetic Spelling: RAY·dee·ay·ter

- Aluminum
Pronunciation link: https://www.merriam-webster.com/dictionary/aluminum
IPA: /əˈluːmənəm/
Phonetic Spelling: uh·LOO·muh·nuhm

- Isoflurane
Pronunciation link: https://www.merriam-webster.com/medical/isoflurane
IPA: /ˌaɪsoʊˈflʊrˌeɪn/
Phonetic Spelling: eye·soh·FLUR·ayn

- Anesthesia
Pronunciation link: https://www.merriam-webster.com/dictionary/anesthesia
IPA: /ˌænəsˈθiːʒə/
Phonetic Spelling: an·uhs·THEE·zhuh

- Prone
Pronunciation link: https://www.merriam-webster.com/dictionary/prone
IPA: /ˈproʊn/
Phonetic Spelling: PROHN

- Supine
Pronunciation link: https://www.merriam-webster.com/dictionary/supine
IPA: /suːˈpaɪn/
Phonetic Spelling: soo·PYN

- Coronal
Pronunciation link: https://www.merriam-webster.com/medical/coronal
IPA: /kəˈroʊnəl/
Phonetic Spelling: kuh·ROH·nuhl

- Sagittal
Pronunciation link: https://www.merriam-webster.com/dictionary/sagittal
IPA: /ˈsædʒətəl/
Phonetic Spelling: SAJ·uh·tuhl

- Axial
Pronunciation link: https://www.merriam-webster.com/dictionary/axial
IPA: /ˈæksiəl/
Phonetic Spelling: AK·see·uhl

- Isocenter
Pronunciation link: No confirmed link found
IPA: /ˈaɪsoʊˌsɛntər/
Phonetic Spelling: EYE·soh·sen·ter

- DICOM
Pronunciation link: No confirmed link found
IPA: /ˈdaɪkɑːm/
Phonetic Spelling: DYE·kom

- Contouring
Pronunciation link: https://www.merriam-webster.com/dictionary/contour
IPA: /ˈkɑːnˌtʊrɪŋ/
Phonetic Spelling: KON·toor·ing

- Interpolation
Pronunciation link: https://www.merriam-webster.com/dictionary/interpolation
IPA: /ɪnˌtɝːpəˈleɪʃən/
Phonetic Spelling: in·ter·puh·LAY·shuhn

- Collimator
Pronunciation link: https://www.merriam-webster.com/dictionary/collimator
IPA: /ˈkɑːləˌmeɪtər/
Phonetic Spelling: KOL·uh·may·ter

- Dose Volume Histogram
Pronunciation link: No confirmed link found
IPA: /doʊs ˈvɑːljuːm ˈhɪstəˌɡræm/
Phonetic Spelling: DOHS VOL·yoom HIS·tuh·gram

- Initialization
Pronunciation link: https://www.merriam-webster.com/dictionary/initialize
IPA: /ɪˌnɪʃələˈzeɪʃən/
Phonetic Spelling: ih·NISH·uh·luh·ZAY·shuhn

- Copper
Pronunciation link: https://www.merriam-webster.com/dictionary/copper
IPA: /ˈkɑːpər/
Phonetic Spelling: KOP·er

- Cone-beam
Pronunciation link: No confirmed link found
IPA: /ˈkoʊnˌbiːm/
Phonetic Spelling: KOHN·beem

- Acquisition
Pronunciation link: https://www.merriam-webster.com/dictionary/acquisition
IPA: /ˌækwəˈzɪʃən/
Phonetic Spelling: ak·wuh·ZISH·uhn

- Segmentation
Pronunciation link: https://www.merriam-webster.com/dictionary/segmentation
IPA: /ˌsɛɡmənˈteɪʃən/
Phonetic Spelling: seg·muhn·TAY·shuhn

- Isodose
Pronunciation link: No confirmed link found
IPA: /ˈaɪsoʊˌdoʊs/
Phonetic Spelling: EYE·soh·dohs

- Superposition
Pronunciation link: https://www.merriam-webster.com/dictionary/superposition
IPA: /ˌsuːpərpəˈzɪʃən/
Phonetic Spelling: soo·per·puh·ZISH·uhn

- Convolution
Pronunciation link: https://www.merriam-webster.com/dictionary/convolution
IPA: /ˌkɑːnvəˈluːʃən/
Phonetic Spelling: kon·vuh·LOO·shuhn

- Monte Carlo
Pronunciation link: https://www.merriam-webster.com/dictionary/Monte%20Carlo
IPA: /ˌmɑːnti ˈkɑːrloʊ/
Phonetic Spelling: MON·tee KAR·loh