# Submission ID #: 71690

# Scriptwriter Name: Sri Lakshmi Priya

# Project Page Link: https://review.jove.com/account/file-uploader?src=21865498

# Title: Workflow for Automated Digital Peripheral Blood and Bone Marrow Cytomorphology, With Consideration of Critical Cellular Misclassifications

# Authors and Affiliations:

# David Starostka1*, Richard Dolezilek2*, Hans Michael Kvasnicka3, Barbora Sotkovska1, Petra Miczkova1

# 1Laboratory of Hemato-oncology and Clinical Biochemistry, Moravian-Silesian Hospital Havirov

# 2Department of Pathology, Moravian-Silesian Hospital Havirov

# 3Institute of Pathology and Molecular Pathology, University Clinic Wuppertal

# *These authors contributed equally

☐   All author names and affiliations are correct (city/state/country information not included in video title page).

# Corresponding Authors:

# David Starostka                        david.starostka@nemhav.cz

# Email Addresses for All Authors:

# David Starostka                        david.starostka@nemhav.cz

# Richard Dolezilek                     richard.dolezilek@nemhav.cz

# Hans Michael Kvasnicka          hansmichael.kvasnicka@helios-gesundheit.de

# Barbora Sotkovska                   barbora.sotkovska@nemhav.cz

# Petra Miczkova                         petra.miczkova@nemhav.cz

## Author Questionnaire

1. Microscopy: Does your protocol require the use of a dissecting or stereomicroscope for performing a complex dissection, microinjection technique, or something similar?  Enter Yes or No.

If Yes, can you record movies/images using your own microscope camera?

Enter Yes or No.

If your protocol involves microscopy but you are not able to record movies/images with your microscope camera, JoVE will need to use our scope kit.

If your microscope does not have a camera port, the scope kit will be attached to one of the eyepieces and you will have to perform the procedure using one eye.

Enter make and model of microscope.

If a dissection or stereo microscope is required for your protocol, please list all shots from the script that will be visualized using the microscope (shots are indicated with the 3-digit numbers, like 2.1.1, 2.1.2, etc.).

Click here to list microscope shots, using the shot numbers from the protocol section of the video script.

2. Software: Does the part of your protocol being filmed include step-by-step descriptions of software usage?  Enter Yes or No.

If Yes, we will need you to record using screen recording software.

We recommend using the screen capture program OBS. JoVE’s tutorial for using OBS Studio is provided at this link: https://review.jove.com/v/5848/screen-capture-instructions-for-authors?status=a7854k

As these files are necessary for finalizing your script, please upload all screen-captured video files to your project page as soon as possible.

3. Filming location: Will the filming need to take place in multiple locations?   Enter Yes or No.

If Yes, how far apart are the locations? Click to enter distance between locations.

4. Testimonials (optional): Would you be open to filming two short testimonial statements live during your JoVE shoot? These will not appear in your JoVE video but may be used in JoVE’s promotional materials. Enter Yes or No.

If Yes, please provide the full name and position (e.g., Director of [Institute Name], Senior Researcher [University Name], etc.) of the author willing to participate.

Enter full author name, Enter author position

To ensure that your script can be filmed in one day, the protocol sections are cumulatively restricted to 55 shots (shots are the 3-digit numbers like 2.1.1, 2.1.2…etc)

Current Protocol Length

Number of Steps:  32

Number of Shots:  58

# Interviews

Videographer: Obtain headshots for all authors available at the filming location.

Answers to these questions will become interview statements that you will deliver on camera.

- Introduction: Answer the 1st REQUIRED prompt and 1 additional question; Conclusion: Answer any 3 questions. No more than 5 interview statements will be included in the video.

- Enter the full name of the author who will deliver the statement.

- Speak naturally and avoid reading the lines.

- Limit the length of each statement to 20 words or fewer.

- Answers will be edited for length, clarity, and consistency with journal style guidelines.

Authors: Please note that the questions will not appear on screen. Please answer in stand-alone full sentence with sufficient context

INTRODUCTION:

REQUIRED: Summarize your research focus and the main questions it aims to answer.

  - Enter author name.: Click here to answer question. Please write in a style that you will be comfortable memorizing and speaking aloud. Limit length to 20 or fewer words.

Describe the research gap or limitation in existing methods and how this protocol addresses them.

  - Enter author name: Click here to answer question. Please write in a style that you will be comfortable memorizing and speaking aloud. Limit length to 20 or fewer words.

Describe the research contexts or systems in which this protocol can be applied.

  - Enter author name: Click here to answer question. Please write in a style that you will be comfortable memorizing and speaking aloud. Limit length to 20 or fewer words.

CONCLUSION:

What does this protocol allow researchers to study or measure?

  - Enter author name: Click here to answer question. Please write in a style that you will be comfortable memorizing and speaking aloud. Limit length to 20 or fewer words.

Describe the most important consideration or challenge when performing this protocol.

  - Enter author name: Click here if you choose this question. Please write in a style that you will be comfortable memorizing and speaking aloud. Limit length to 20 or fewer words.

Are there any additional methods or analyses that can be performed following this procedure?

  - Enter author name: Click here if you choose this question. Please write in a style that you will be comfortable memorizing and speaking aloud. Limit length to 20 or fewer words.

How can future studies build on this work? Describe the questions or research directions that can be investigated next.

  - Enter author name: Click here if you choose this question. Please write in a style that you will be comfortable memorizing and speaking aloud. Limit length to 20 or fewer words.

Videographer: Obtain headshots for all authors available at the filming location.

# Ethics Title Card

This research has been approved by the Local Ethics Committee at Hospital Havířov. Written consent was obtained from all the participants

# Protocol

Please review this section to make sure that it accurately describes your protocol. Use Track Changes when making edits or revisions.

- The two-digit steps (e.g., 2.1., 2.2.) with purple font are the narration.  JoVE is responsible for the narration of the protocol and results.

- Red italics are pronunciation guides indicating how the word will be spoken.

- Filming should take no more than 10 minutes per step. If a step takes more than 10 minutes, prepare the product for that step in advance.

- The three-digit shots (e.g., 2.1.1., 2.2.2.) are the actions that the videographer will capture.

Videographer: Please always record an audio slate while filming each shot mentioned in the shotlist.

- Automated Peripheral Blood Smear Preparation and Staining

Demonstrator: Click here to enter name of demonstrator(s)

If the same person is the demonstrator throughout, mention them once here and remove the "Demonstrator" field from the other sections; if the demonstrator changes, retain the field in the respective sections.

  - To begin, check the tubing condition [1], network connection [2], and waste container status [3]. Confirm that sufficient reagents and an adequate supply of microscope slides are available [4].

    - WIDE: Talent inspecting the tubing condition. Videographer: Please obtain CU shot as well for this shot

    - Talent checking the network connection.

    - Talent inspecting the waste container status.

    - Talent arranging the reagents and microscope slides on the workbench.

  - Turn on the instrument [1]. Allow the main unit to perform the automatic self-check [2].

    - Talent turning on the instrument.

    - SCREEN: Show the instrument performing the automatic self-check.

Authors: Please create screen capture videos of the shots labeled as SCREEN,

create a screenshot summary, and upload the files to your project page as soon as possible:  https://review.jove.com/account/file-uploader?src=21865498

Videographer: Please film the computer screen as a back up

  - Log in to the software using a valid username and password and operate the instrument using the current software versions [1].

    - SCREEN: Enter a valid username and password and log in to the software, and show the software home screen.

  - Select operations from the Menu [1]. Manage requests for slide preparation and smear preparation in the Worklist [2]. Review request data in the Browser [3]. Check the smear preparation status in the Status tab [4].

    - SCREEN: Open the Menu and select the operations.

    - SCREEN: Open the Worklist and display requests for slide preparation and smear preparation.

    - SCREEN: Open the Browser and review the request data.

    - SCREEN: Open the Status tab and display the smear preparation status.

  - Download a new slide preparation request from the host computer or create a new request [1].

    - SCREEN: Download a new slide preparation request from the host computer or create a new request.

  - Edit or delete slide preparation requests as needed [1].

    - SCREEN: Edit or delete a slide preparation request.

  - Wear latex or nitrile gloves before handling peripheral blood and bone marrow samples and smears [1]. Select the smear preparation mode [2]. Place the sample rack into the sampler [3]. Check that the spreader glass is intact, completely clean, and securely placed in the holder [4]. Start smear preparation and staining from the dialog window [5].

    - Talent putting on latex or nitrile gloves.

    - SCREEN: Select the smear preparation mode.

    - Talent placing the sample rack into the sampler.

    - Talent inspecting the spreader glass in the holder.

    - SCREEN: Smear preparation and staining are being started from the dialog window.

  - Check the status of the smear being prepared and review the details of the slide preparation requests [1].

    - SCREEN: Display the smear preparation status and the details of the slide preparation requests.

  - Remove the sample rack [1] and the prepared and stained slides [2].

    - Talent removing the sample rack.

    - Talent removing the prepared and stained slides.

  - After performing the instrument inspection and cleaning, turn off the instrument [1].

    - Talent turning off the instrument.

- Automated Peripheral Blood Smear Analysis

Demonstrator: Click here to enter name of demonstrator(s)

  - Use the application for the differential count of white blood cells, assessment of red blood cell morphology, and estimation of platelet count [1]. Verify the proposed classification of each cell type [2].

    - SCREEN: Display the application showing the white blood cell differential count, red blood cell morphology assessment, and platelet count estimation.

    - SCREEN: Show the proposed classification of each cell type.

  - Once the analyzer is started, log in to the software using a valid username and password and operate the instrument using the current software versions [1].

    - SCREEN: Enter a valid username and password and log in to the software, and show the software home screen.

  - Open the System Control View tab for the quality control procedure and processing orders [1].

    - SCREEN: Open the System Control View tab.

  - Perform the daily internal quality control using a control slide [1]. Run the cell detection test once daily to detect at least 100 cells [2].

    - Talent placing the control slide into the analyzer.

    - SCREEN: Start the cell detection test, recognizing 100 cells.

  - Prepare the peripheral blood smear with the slide-making and staining instrument [1]. Place the stained smear into the analyzer loading drawer [2].

    - Talent holding the prepared peripheral blood smear.

    - Talent placing the stained smear into the analyzer loading drawer.

  - Enter the test identification manually or with a barcode scanner [1]. Allow the request from the laboratory information system to load automatically [2]. Close the analyzer input door [3].

    - SCREEN: Enter the test identification manually or scan the barcode to display the test identification.

    - SCREEN: Show the laboratory information system request loading automatically.

    - Talent closing the analyzer input door.

  - Process the smear [1].

    - SCREEN: Show smear processing being started.

  - Review the preliminary results [1]. Reclassify cells as necessary on the analyzer or using the remote review software, considering the three proposed classification options and their corresponding probability percentages [2].

    - SCREEN: Display the preliminary analysis results.

    - SCREEN: Reclassify a cell using the analyzer or remote review software while displaying the proposed classification options and probability percentages.

  - In WBC mode, review the preliminary classification of nucleated cells [1]. Reassign the cell by dragging its image to the correct category if an incorrect classification is identified [2].

    - SCREEN: Display the nucleated cell classification in WBC mode.

    - SCREEN: Drag a cell image to the correct classification category.

  - In RBC mode, review the assessment of RBC morphological changes and correct inaccurate evaluations or add additional pathological features [1].

    - SCREEN: Display the red blood cell morphology being assessed in RBC mode and correct the evaluation or select additional pathological features.

  - In platelet or PLT mode, review the platelet evaluation [1].

    - SCREEN: Display the platelet evaluation in PLT mode.

  - Add comments manually or select predefined text options [1]. Confirm the final result so it can be sent automatically to the laboratory information system [2].

    - SCREEN: Enter comments manually or select predefined text options.

    - SCREEN: Confirm the final result and display successful transmission to the laboratory information system.

  - After logging out of the software, shut down the analyzer [1].

    - Talent shutting down the analyzer.

- Automated Bone Marrow Differential Cell Analysis

Demonstrator: Click here to enter name of demonstrator(s)

  - Use the system for the differential count of PB and BM cells and verify the proposed classification of each cell type [1].

    - SCREEN: Display the system showing peripheral blood and bone marrow cell differential count results and show the proposed classification of each cell type.

  - After starting the analyzer, log in to the software using a valid username and password and operate the instrument using the current software versions [1].

    - SCREEN: Enter a valid username and password and log in to the software, and show the software home screen.

  - Use Project for processing orders and slide management [1]. Use Edit to configure classification rules and the pre-classification strategy [2]. Use Report for report formatting and generation [3]. Use the Manage tab for account settings and signature rules [4].

    - SCREEN: Open the Project tab and show order processing and slide management being performed.

    - SCREEN: Open Edit and display the configuration of the classification rule and pre-classification strategy.

    - SCREEN: Open Report and demonstrate report format settings and report generation.

    - SCREEN: Open the Manage tab and demonstrate account settings and signature rule configuration.

  - Use a BM smear stained with the staining instrument [1]. Enter the test identification manually or with a barcode or QR code scanner [2]. Allow the request from the laboratory information system to load automatically [3]. Place the stained smears into the analyzer loading drawer and close the analyzer input door [4].

    - Talent holding the stained bone marrow smear.

    - SCREEN: Enter the test identification manually or scan the barcode or QR code.

    - SCREEN: Show the laboratory information system request loading automatically.

    - Talent placing the stained smears into the analyzer loading drawer and closing the analyzer input door.

  - Process the smear [1].

    - SCREEN: Show the smear processing being started.

  - After reviewing the preliminary results, reclassify cells where necessary on the analyzer or using the remote review software, and reassign incorrectly classified cells to the correct category [1].

    - SCREEN: Show the cells being reclassified on the analyzer or using the remote review software.

  - Add comments and cytomorphological descriptions manually or by selecting predefined text options [1].

    - SCREEN: Show the comments and cytomorphological descriptions being added manually or by selecting predefined text options.

  - Once the preliminary report is completed, confirm the final result so it can be sent automatically to the laboratory information system [1].

    - SCREEN: Show the final result being confirmed and display the successful transmission to the laboratory information system.

  - After logging out of the software, shut down the analyzer [1].

    - Talent shutting down the analyzer.

# Results

Please review this section to make sure that it accurately reflects your findings.

- This section will not be recorded by the videographer. It only includes the figures/tables from your manuscript (called LAB MEDIA).

- Use Track Changes when making edits or revisions. Ensure the voiceover length is below 200 words. Current word count: 120.

- Please note that the video cannot include voiceover without an accompanying visual.

- Results

  - Correctly classified cells generally displayed typical cytomorphological features [1], whereas incorrectly classified cells frequently showed atypical morphology [2].

    - LAB MEDIA: Figure 4A.

    - LAB MEDIA: Figure 5.

  - Misclassifications most often occurred between atypical lymphocytes and blasts [1], myeloblasts and lymphocytes [2], lymphoblasts and lymphocytes [3], monoblasts or promonocytes and promyelocytes [4], and immature plasma cells and blasts [5].

    - LAB MEDIA: Figure 4B. Video editor: In the middle panel, highlight the rows where Atypical lymphocyte appears under Expert classification and Blast, NOC under ADM pre-classification.

    - LAB MEDIA: Figure 5. Video editor: Highlight the 2 rows with AML under diagnosis.

    - LAB MEDIA: Figure 5. Video editor: Highlight the rows with Lymphoblast under Expert classification and Lymphocyte under ADM classification.

    - LAB MEDIA: Figure 5. Video editor: Highlight the examples with Monoblast under Expert classification and Promyelocyte and Blast, NOC under ADM classification.

    - LAB MEDIA: Figure 5. Video editor: Highlight the row with Immature plasma cell under Expert classification and Blast, NOC under ADM classification.

  - In lymphocytic neoplasms, representative examples of misclassified neoplastic lymphocytes were identified [1].

    - LAB MEDIA: Figure 5. Video editor: In the left panel, highlight the rows labelled Splenic MZL, MCL, HCL, MZL, and CLL, pleomorphic variant.

  - In acute leukemia and chronic myelomonocytic leukemia, representative examples showed misclassified lymphoblasts, myeloblasts, monoblasts, and promonocytes [1].

    - LAB MEDIA: Figure 5. Video editor: In the right panel, highlight the rows labelled T-ALL and AML.

  - In multiple myeloma, misclassifications occurred within the plasma cell lineage, particularly involving plasmablasts [1] and immature plasma cells [2].

    - LAB MEDIA: Figure 5. Video editor: In the right panel, highlight the rows labelled Plasmablast under Expert Classification.

    - LAB MEDIA: Figure 5. Video editor: In the bottom right panel, highlight the rows labelled Immature plasma cell under Expert Classification.

  - Critical misclassifications of neoplastic lymphocytes [1], small lymphoblasts or myeloblasts [2], granular monoblasts may pose a risk of misdiagnosis [3].

    - LAB MEDIA: Figure 5. Video editor: Highlight the rows showing atypical lymphocyte under Expert Classification.

    - LAB MEDIA: Figure 5. Video editor: Highlight the rows showing lymphoblast and myeloblast under Expert Classification.

    - LAB MEDIA: Figure 5. Video editor: Highlight the row showing the Monoblast under Expert Classification.

  - Similar risks were associated with critical misclassifications of dysplastic promonocytes and monocytes [1], plasmablasts [2], and immature plasma cells [3].

    - LAB MEDIA: Figure 4. Video editor: Highlight the rows showing Promonocyte and Monocyte under ADM pre-classification.

    - LAB MEDIA: Figure 5. Video editor: Highlight the rows showing Plasmablast under Expert Classification.

    - LAB MEDIA: Figure 5. Video editor: Highlight the rows showing Immature plasma cell under Expert Classification.