# Submission ID #: 72062

# Scriptwriter Name: Pallavi Sharma

# Project Page Link: https://review.jove.com/files_upload.php?src=21978958

# Title: Workflows for the Photopic Negative Response and the Baby Vision Test

# Authors and Affiliations:

Tingting Yan1, Xu Chen1, Guanshun Yu1

1National Clinical Research Center for Ocular Diseases, Eye Hospital, Wenzhou Medical University

# Corresponding Authors:

Xu Chen					(kellychen106@eye.ac.cn)

Guanshun Yu					(ygs@eye.ac.cn)

# Email Addresses for All Authors:

Tingting Yan 					(yantt@eye.ac.cn)

Xu Chen					(kellychen106@eye.ac.cn)

Guanshun Yu					(ygs@eye.ac.cn)

## Author Questionnaire

1. We have marked your project as author-provided footage, meaning you film the video yourself and provide JoVE with the footage to edit. JoVE will not send the videographer. Please confirm that this is correct.

√ Correct

2. Microscopy: Does your protocol require the use of a dissecting or stereomicroscope for performing a complex dissection, microinjection technique, or something similar?  No.

3. Software: Does the part of your protocol being filmed include step-by-step descriptions of software usage?  Yes.

Authors: Please create screen capture videos of the shots labeled as SCREEN, create a screenshot summary, and upload the files to your project page as soon as possible: https://review.jove.com/files_upload.php?src=21978958

4. Proposed filming date: To help JoVE process and publish your video in a timely manner, please indicate the proposed date that your group will film here: August 3–9, 2026.

When you are ready to submit your video files, please contact our Content Manager, Utkarsh Khare. st

Current Protocol Length

Number of Steps:  21

Number of Shots:  42

# Interviews

Authors: Please note that the questions will not appear on screen. Please answer in stand-alone full sentence with sufficient context.

INTRODUCTION:

REQUIRED: Summarize your research focus and the main questions it aims to answer.

  - Guanshun Yu: Our research focuses on standardizing PhNR and Baby Vision workflows to improve their clinical utility for diagnosing and monitoring ocular diseases.

    - INTERVIEW: Named Talent says the statement above in an interview-style shot, looking slightly off-camera.

Describe the research contexts or systems in which this protocol can be applied.

  - Guanshun Yu: This protocol is applicable to children's eye care, electrophysiology laboratories, and routine clinics for visual disorders.

    - INTERVIEW: Named Talent says the statement above in an interview-style shot, looking slightly off-camera.

CONCLUSION:

What does this protocol allow researchers to study or measure?

  - Guanshun Yu: This protocol allows researchers to measure retinal ganglion cell function and visual acuity thresholds in young children.

    - INTERVIEW: Named Talent says the statement above in an interview-style shot, looking slightly off-camera.

Ethics Title Card

This research has been approved by the Institutional Review Board (IRB) at Eye Hospital, Wenzhou Medical University

# Protocol

- Equipment Preparation for Photopic Negative Response (PhNR)

Demonstrators: Mengya Huai, Yongjie Hu

  - To begin, perform the routine calibration and verification of all stimuli and recording instruments according to the ISCEV (I-S-C-E-V) recommendations [1]. Verify that the electroretinography system is properly connected to the visual stimulator [2].

    - Talent checking the stimulus and recording instruments.

    - Talent checking the connection between the electroretinography system and the visual stimulator.

  - Inspect all cable connections and confirm that they are intact and functioning correctly [1].

    - Talent inspecting all cable connections.

  - Using alcohol-soaked cotton swabs, disinfect the headrest and the chinrest before testing [1].

    - Talent disinfecting the headrest and the chinrest with alcohol-soaked cotton swabs.

  - Next, turn on the computer and wait until the desktop is fully loaded [1]. Switch the stimulator rocker switch from O to I [1].

    - Talent turning on the computer.

    - Talent switching the stimulator rocker switch from O to I.

  - Launch the acquisition software [1], select the Examination menu, and choose the  PhNR (P-H-N-R) Protocol [2].

    - SCREEN: The acquisition software being launched.

Authors: Please create screen capture videos of the shots labeled as SCREEN, create a screenshot summary, and upload the files to your project page as soon as possible: https://review.jove.com/files_upload.php?src=21978958

    - SCREEN: Examination menu being opened. PhNR Protocol being selected.

  - Next, ensure that the camera is functioning properly [1] .

    - Talent checking that the camera is functioning correctly.

- Patient Preparation and Pupil Dilation for PhNR Recording

  - Explain the procedure to the parent or guardian, emphasizing that the examination is noninvasive and painless [1]. Confirm that the child is awake and calm. Avoid testing children who are drowsy, dozing, or actively crying [2].

    - Talent explaining the procedure to the parent or guardian.

    - Talent confirming that the child is awake and calm.

  - Confirm the baseline pupil status [1]. Then, instill 1 drop of 0.5 percent tropicamide every 3 to 5 minutes for a total of 3 applications [2].

    - Talent observing the pupil under normal room light.

    - Talent instilling 1 drop of 0.5 percent tropicamide into the eye.

  - Confirm adequate pupil dilation after 30 minutes [1]. Allow the patient to adapt to photopic conditions for at least 10 minutes before recording [1]

    - Talent checking the pupil dilation.

    - Talent positioning the patient under photopic conditions before recording.

- Electrode Preparation and Placement for PhNR Recording

  - Inspect all skin electrodes and confirm that they are intact, uncontaminated, and free of deterioration [1]. Apply gel to the cleaned periocular skin [2].

    - Talent inspecting all skin electrodes.

    - Talent applying gel to the cleaned periocular skin.

  - Fill the skin electrodes with conductive paste [1] and secure them with medical tape [2].

    - Talent filling the skin electrodes with conductive paste.

    - Talent securing the electrodes with medical tape.

  - Place the active electrodes on the lower eyelids [1], the reference electrodes at the external canthi [2], and the ground electrode on the forehead [3]. Measure electrode impedance and reapply electrodes if it exceeds the acceptable range [2].

    - Talent placing the active electrodes on the lower eyelids.

    - Talent placing the reference electrodes at the external canthi.

    - Talent placing the ground electrode on the forehead.

    - SCREEN: Electrode impedance is being measured.

- Photopic Negative Response Examination

  - Click the Patient ID or Eye icon, then enter the patient's information [1] and select binocular testing [2].

    - SCREEN: Patient ID/Eye icon being clicked. Patient information being entered.

    - SCREEN: Binocular testing being selected.

  - Position the child comfortably at the examination station [1]. Adjust the table height [1] and position the chinrest [2].

    - Talent positioning the child comfortably at the examination station.

    - Talent adjusting the table height.

    - Talent positioning the chinrest.

  - Recheck the electrode impedance [1] and click on Exam to begin the acquisition [1].

    - SCREEN: Electrode impedance values being rechecked.

    - SCREEN: Exam being clicked to begin the acquisition.

  - Monitor the waveform and the artifact-rejection indicator continuously [1].

    - SCREEN: The waveform and the artifact-rejection indicator being monitored continuously.

  - Instruct the patient to fixate on the central fixation target and encourage the patient to minimize blinking during the acquisition [1]. Use verbal cues to maintain the patient's attention when cooperation is limited [2].

    - Talent instructing the patient to fixate on the central fixation target.

    - Talent encouraging the patient to minimize blinking during the acquisition.

    - Talent using verbal cues to maintain the patient's attention.

  - Continue recording until a reliable waveform is obtained [1], then stop the acquisition by selecting Results [1].

    - SCREEN: Waveform recording continuing until a reliable waveform is obtained.

    - SCREEN: Results being clicked to stop the acquisition.

  - Next, remove all electrodes [1]. Using a cotton swab, clean the skin [2]. Using alcohol-soaked cotton swabs, disinfect the headrest and the chinrest [3].

    - Talent removing all electrodes.

    - Talent cleaning the skin with a cotton swab.

    - Talent disinfecting the headrest and the chinrest with alcohol-soaked cotton swabs.

  - Display the averaged waveform [1] and identify the a-wave, the b-wave, and the photopic negative response trough [2].

    - SCREEN: The averaged waveform being displayed.

    - SCREEN: The a-wave, the b-wave, and the photopic negative response trough being identified on the averaged waveform.

  - Measure the photopic negative response amplitude from the baseline to the trough of the negative wave following the b-wave, and the peak time from stimulus onset to the PhNR trough [1].

    - SCREEN: The photopic negative response amplitude being measured. The peak time being measured.

# Results

- Results

  - Following implementation of the 6E protocol, the PhNR success rate increased from 61.54% in the year before implementation [1] to 86.45% in the year after implementation [2].

    - LAB MEDIA: Figure 8. Video editor: Highlight the blue bar above the label 'One year prior'.

    - LAB MEDIA: Figure 8. Video editor: Highlight the blue bar above the label 'One year later'.

  - The number of valid Baby Vision examination results increased substantially during the 2-year period before implementation [1] and the 2-year period after implementation [2].

    - LAB MEDIA: Figure 9. Video editor: Highlight the blue bar labelled 'valid recordings' above 'Two years prior'.

    - LAB MEDIA: Figure 9. Video editor: Highlight the blue bar labelled 'valid recordings' above 'Two years later'.

  - The mean Baby Vision examination duration decreased from 17.44 minutes during the 2-year period before implementation [1] to 13.28 minutes during the 2-year period after implementation [2].

    - LAB MEDIA: Figure 9. Video editor: Highlight the orange bar labelled 'average test duration' above 'Two years prior'.

    - LAB MEDIA: Figure 9. Video editor: Highlight the orange bar labelled 'average test duration' above 'Two years later'.