# Submission ID #: 73445

# Scriptwriter Name: Vipasha Sharma

# Project Page Link: https://review.jove.com/account/file-uploader?src=22400773

# Title: Fabrication and Use of a 32-Well LED-Embedded Microplate for Optogenetic Dynamic Control

# Authors and Affiliations:

# Bhavya Jaiswal¹, Trevor S. Black¹, Hari R. Namboothiri1, Krishna Pochana2, Chelsea Y. Hu¹

¹Department of Chemical Engineering, Texas A&M University

2Department of Electrical Engineering and Computer Science, Massachusetts Institute of Technology

☒   All author names and affiliations are correct (city/state/country information not included in video title page).

# Corresponding Authors:

Chelsea Y. Hu			chelsea.hu@tamu.edu

# Email Addresses for All Authors:

bav45@tamu.edu

trevor.black@tamu.edu

harikrishnan@tamu.edu

kpochana@mit.edu

chelsea.hu@tamu.edu

## Author Questionnaire

1. Microscopy: Does your protocol require the use of a dissecting or stereomicroscope for performing a complex dissection, microinjection technique, or something similar?  No

2. Software: Does the part of your protocol being filmed include step-by-step descriptions of software usage?  Yes

If Yes, we will need you to record using screen recording software.

We recommend using the screen capture program OBS. JoVE’s tutorial for using OBS Studio is provided at this link: https://review.jove.com/v/5848/screen-capture-instructions-for-authors?status=a7854k

As these files are necessary for finalizing your script, please upload all screen-captured video files to your project page as soon as possible.

3. Filming location: Will the filming need to take place in multiple locations?   No

4. Testimonials (optional): Would you be open to filming two short testimonial statements live during your JoVE shoot? These will not appear in your JoVE video but may be used in JoVE’s promotional materials. No

To ensure that your script can be filmed in one day, the protocol sections are cumulatively restricted to 55 shots (shots are the 3-digit numbers like 2.1.1, 2.1.2…etc)

Current Protocol Length

Number of Steps: 20

Number of Shots: 59

# Interviews

Videographer: Obtain headshots for all authors available at the filming location.

Answers to these questions will become interview statements that you will deliver on camera.

- Introduction: Answer the 1st REQUIRED prompt and 1 additional question; Conclusion: Answer any 3 questions. No more than 5 interview statements will be included in the video.

- Enter the full name of the author who will deliver the statement.

- Speak naturally and avoid reading the lines.

- Limit the length of each statement to 20 words or fewer.

- Answers will be edited for length, clarity, and consistency with journal style guidelines.

Authors: Please note that the questions will not appear on screen. Please answer in stand-alone full sentence with sufficient context

INTRODUCTION:

REQUIRED: Summarize your research focus and the main questions it aims to answer.

  - Bhavya: Our work involves developing an accessible platform to study how dynamic light inputs control gene expression in microbial cultures.

Describe the research gap or limitation in existing methods and how this protocol addresses them.

  - Bhavya: Existing optogenetic platforms are static, complex, or sensor-limited. LEMOS 2.0 enables dynamic stimulation with accurate measurements in a standard microplate reader.

CONCLUSION:

What does this protocol allow researchers to study or measure?

  - Trevor: This protocol allows researchers to deliver optogenetic stimulation and measure gene expression and growth dynamics in microbial cultures simultaneously, within a standard microplate reader.

Describe the most important consideration or challenge when performing this protocol.

  - Trevor: A key consideration is avoiding bubbles or debris in the PDMS microwells, which can interfere with accurate optical measurements.

How can future studies build on this work? Describe the questions or research directions that can be investigated next.

  - Bhavya: The demonstrated framework can be extended to other microbial strains, fluorescent reporters, optogenetic systems and more advanced control algorithms.

Videographer: Obtain headshots for all authors available at the filming location.

# Ethics Title Card

This research has been approved by the Texas A&M Institutional Biosafety Committee (IBC).

# Protocol

Please review this section to make sure that it accurately describes your protocol. Use Track Changes when making edits or revisions.

- The two-digit steps (e.g., 2.1., 2.2.) with purple font are the narration.  JoVE is responsible for the narration of the protocol and results.

- Red italics are pronunciation guides indicating how the word will be spoken.

- Filming should take no more than 10 minutes per step. If a step takes more than 10 minutes, prepare the product for that step in advance.

- The three-digit shots (e.g., 2.1.1., 2.2.2.) are the actions that the videographer will capture.

Videographer: Please always record an audio slate while filming each shot mentioned in the shotlist.

- PDMS Microwell Casting

Demonstrator: Bhavya

If the same person is the demonstrator throughout, mention them once here and remove the "Demonstrator" field from the other sections; if the demonstrator changes, retain the field in the respective sections.

  - To begin, arrange the materials in the fume hood [1]. Sanitize the interior of the LEMOS 2.0 (Lee-moss) frame where the microwells will be cast by spraying 70% ethanol, avoiding direct contact with any electrical components except the light-emitting diodes, or LEDs (L-E-D). Let the frame dry completely before casting the Polydimethylsiloxane, or PDMS (P-D-M-S) microwells [2].

    - WIDE: Talent arranging the material

    - Talent sanitizing the interior of LEMOS frame by spraying 70% ethanol and keeping it aside for drying

  - Spray the PDMS microwell casting mold with a release agent and let it settle on the plastic surface for 10 minutes [1-TXT].

    - Talent spraying the mold with release agent and keeping it aside  TXT: Release agent: Ease Release 200

  - Prepare the PDMS mixture by weighing 25 grams of the 184-silicone elastomer base and 2.5 grams of the curing agent [1]. Mix thoroughly in a disposable plastic cup [2-TXT].

    - Talent weighing 25 g of base and 2.5 g of curing agent

    - Talent mixing the base and curing agent TXT-Base: Curing agent = 10: 1 (w/w)

  - Mix the PDMS vigorously for at least 5 minutes using a disposable spatula or wooden stick, scraping the sides and bottom of the cup to ensure thorough mixing [1]. Place the cup containing the PDMS mixture in a vacuum desiccator [2] and set the vacuum to at least 65 cm Mercury for 45-60 minutes, or until all visible air bubbles have been removed [3].

    - Talent mixing the PDMS mixture vigorously using a wooden stick

    - Talent placing the cup in a vacuum desiccator

    - Talent setting the vacuum to 65 cm Hg  TXT: Use a container ≥3× the mixture volume to prevent overflow

  - Secure the LEMOS 2.0 frame to the outer shell using appropriately sized binder clips on all four sides [1]. Carefully fill each microwell with the degassed PDMS using a syringe, while minimizing the formation of air bubbles [2] [3].

    - Talent securing the LEMOS 2.0 frame to the outer shell using binder clips

    - Talent carefully pouring the degassed PDMS into the frame

    - CU: Talent using a syringe to fill the microwells with PDMS

  - Insert the 3D-printed casting mold into the PDMS-filled frame [1]. Gently press the mold into place, ensuring the circular protrusions align with the holes in the LEMOS 2.0 frame [2]. Wipe away any excess PDMS from the edges using a clean, lint-free wipe [3]. Let the PDMS cure on a level surface at room temperature for at least 30 hours [4].

    - Talent inserting the casting mold into the PDMS-filled frame

    - Talent gently pressing the mold

    - Talent wiping the excess PDMS using a clean, lint-free wipe

    - Talent placing the device on a level surface for curing

  - After curing, remove the casting mold by gripping the edges and lifting it with a steady force, ensuring the PDMS microwells remain in the frame [1]. Inspect each microwell to ensure that the walls are optically clear and the bottom is free of air bubbles [2].

    - Talent removing the casting mold

    - Talent inspecting the microwells

- Microplate Reader Configuration and Culture Preparation

Demonstrator: Bhavya

  - Open the closed-loop protocol experimental file from the GitHub repository in Gen5 (Gen-five) software [1]. Open the Plate Types menu and import the LEMOS 2.0 plate layout from GitHub [2].

    - SCREEN: Closed loop protocol file being opened in Gen5 software

    - SCREEN: Plate Types menu being opened and the LEMOS 2.0 plate layout being imported

Authors: Please create a screen capture video of the shots labelled SCREEN and upload them onto your project page: https://review.jove.com/account/file-uploader?src=22400773

  - Add 200 µL of sterile M9CA medium to each microwell [1] and seal the LEMOS 2.0 device with a medical-grade gas-permeable sealing film [2]. Perform a baseline measurement by recording the OD₆₀₀ (O-D- six hundred) and fluorescence values for all 32 microwells [3] and compare the measurements with those obtained from a standard 96-well microplate [4].

    - Talent adding the sterile medium to each well

    - Talent sealing the device with sealing film

    - SCREEN: OD₆₀₀ and Fluorescence baseline measurements being acquired for all 32 microwells

    - SCREEN: Comparison of the baseline measurements with those from a standard 96-well microplate

  - Inoculate a single colony of the CcaSR (C-c-a-S-R) optogenetic strain and fluorescence-negative control strain into separate 14-milliliter round-bottom culture tubes containing 2 milliliters of M9CA broth supplemented with 50 microgram/milliliter spectinomycin and 25 microgram/milliliter chloramphenicol [1]. Cover both tubes with aluminum foil [2] and incubate overnight at 37 degrees Celsius with shaking at 220 rotations per minute [3].

    - Talent transferring the optogenetic and negative-control strains into separate culture tubes

    - Talent covering the tubes with the aluminum foil

    - Talent placing the tubes at 37 °C with shaking at 220 rpm

  - Transfer 100 microliters of each overnight culture into 2 milliliters of fresh medium [1] and incubate for an additional 4 hours at 37 degrees Celsius with shaking at 220 rotations per minute [2]. After 4 hours, measure the OD₆₀₀ of each subculture using a microplate reader [3-TXT].

    - Talent transferring 100 µL of each overnight culture into fresh medium

    - Talent placing the tubes at 37 °C, 220 rpm

    - SCREEN: OD₆₀₀ measurement being recorded  TXT: Expected OD₆₀₀: 0.4–0.6

  - Dilute each culture to an OD₆₀₀ of 0.1 in fresh medium [1], dispense 200 microliters into each microwell of the LEMOS 2.0 device, avoiding air bubbles [2], and seal the device with a medical-grade sealing film [3].

    - Talent diluting the culture using fresh medium

    - Talent adding 200 µL into each well

    - Talent sealing the device with a sealing film

- Optogenetic Time-Course Experiment

Demonstrator: Bhavya

  - Turn on the LEMOS 2.0 device by sliding the power switch to the left [1], as indicated by the illumination of the first green LED on the LED strip [2]. Connect the central Arduino to the computer via USB [3] and select it from the drop-down menu in the Arduino IDE [4].

    - Talent switching on the LEMOS 2.0 device.

    - CU: Shot of the first LED on the LED strip illuminating green

    - Talent connecting the central Arduino to the computer via USB.

    - SCREEN: Click the Arduino IDE dropdown, select the Arduino board

  - Open the Serial Monitor [1] and connect the peripheral Arduino in the LEMOS 2.0 device to the central Arduino [2].

    - SCREEN: Serial Monitor being opened and the connection message is displayed on screen

    - Talent connecting the peripheral Arduino in the LEMOS 2.0 device to the central Arduino

  - Turn off the indicator LED by entering the letter “o” in the Serial Monitor [1] [2]. Open the LEMOS-start-data experiment file in Gen5 [3]. Record the baseline OD₆₀₀ and fluorescence values of the cultures [4]. Confirm that the measurements are within the expected range, then close the experiment file [5].

    - SCREEN: Letter "o" being entered in the serial monitor

    - CU: Shot of Indicator LED turning off

    - SCREEN: Opening the LEMOS-start-data.xpt file in Gen5

    - SCREEN: Baseline OD₆₀₀ and fluorescence readings being acquired

    - SCREEN: Experiment file being closed after confirming the measurements range

  - Open the closed loop protocol file from the GitHub repository in Gen5 and save it as an experiment [1]. Open the control duty cycle script for duty cycle experiments or the control pid.py script for PID experiments in a compatible Python IDE [2].

    - SCREEN: Closed loop protocol file being opened and saved

    - SCREEN: Control duty cycle script in the IDE being opened

  - In the Main section of the Python script, modify the file variable to match the filename of the Gen5 experiment [1]. Set the Gen5 experiment and the Python script side by side on the screen [2]. Start the Gen5 experiment and abort it once the kinetic step begins [3]. Run the Python script while keeping the Gen5 window open and unobstructed [4].

    - SCREEN: File variable being modified

    - SCREEN: Gen5 experiment and the Python script being set side by side

    - SCREEN: Gen5 experiment being started and aborted once the kinetic step begins

    - SCREEN: Python script being run

  - Import the exported plate-reader data into the analysis notebook [1]. Calculate normalized fluorescence to generate time-course plots [2].

    - SCREEN: Data being imported into the analysis

    - SCREEN: Normalized fluorescence being calculated and plots being generated

- Device cleaning

Demonstrator: Bhavya

  - Gently remove and discard the medical-grade sealing film from the LEMOS 2.0 device [1]. Using a pipette, aspirate and discard the cell suspension from each microwell according to biological waste disposal procedures [2]. Separate the outer shell from the LEMOS 2.0 frame using a steady, even force, taking care not to bend or twist the PLA frame excessively [3].

    - Talent removing and discarding the sealing film

    - Talent aspirating and discarding the cell suspension

    - Talent separating the outer shell from the LEMOS 2.0 frame

  - Peel away the bulk PDMS layer from the frame [1]. Then, use tweezers to carefully remove the PDMS microwells from the LEMOS 2.0 frame [2]. Sanitize the LEMOS 2.0 frame and outer shell with 70% ethanol, avoiding direct contact with the electrical components except the LEDs [3].

    - Talent peeling away the bulk PDMS layer from the frame

    - Talent removing the PDMS microwells using the tweezers

    - Talent spraying 70% ethanol onto the LEMOS 2.0 frame and outer shell

# Results

Please review this section to make sure that it accurately reflects your findings.

- This section will not be recorded by the videographer. It only includes the figures/tables from your manuscript (called LAB MEDIA).

- Use Track Changes when making edits or revisions. Ensure the voiceover length is below 200 words. Current word count: 142

- Please note that the video cannot include voiceover without an accompanying visual.

- Results

  - Cells exposed to constant green light showed a marked increase in fluorescence over time [1]. In contrast, dark microwells adjacent to the illuminated wells exhibited fluorescence levels comparable to those of isolated dark controls, indicating minimal light leakage or inter-well crosstalk [2].

    - LAB MEDIA: Figure 3D   Video Editor: Please highlight the green curve and its corresponding green legend in the left graph

    - LAB MEDIA: Figure 3D   Video Editor: Please sequentially highlight the yellow, purple, and grey curves together, along with their corresponding legend in the left graph

  - Increasing the green-light duty cycle produced progressively higher expression of superfolder green fluorescent protein over time, demonstrating that gene expression can be tuned by adjusting the duty cycle [1]. The expression rate increased with duty cycle up to 40%, after which it reached a plateau, indicating saturation at higher green-light stimulation [2].

    - LAB MEDIA: Figure 3E   Video Editor: Please sequentially highlight the fluorescence curves from 5% to 80% duty cycle

    - LAB MEDIA: Figure 3F   Video Editor: Please highlight the data points from 5% to 40%, followed by the 40%, 60%, and 80% points together

  - The orange set-point tracking curve gradually approached the predefined set point [1] and remained close to it [2], while the constant green and red-light conditions remained above and below the target, respectively [3]. Correspondingly, the green-light duty cycle progressively decreased as the fluorescence signal approached the set point [4] and remained low after the target was reached [5].

    - LAB MEDIA: Figure 4B  Video Editor: Please highlight the solid orange set-point tracking curve

    - LAB MEDIA: Figure 4B  Video Editor: Please highlight the dotted orange set-point line and the orange tracking curve together around 7-12 hours

    - LAB MEDIA: Figure 4B  Video Editor: Please highlight the orange dotted set-point line, followed by the green constant-light curve and the red constant-light curve

    - LAB MEDIA: Figure 4C  Video Editor: Please highlight the green duty-cycle bars from approximately 3 to 7 h

    - LAB MEDIA: Figure 4C  Video Editor: Please highlight the green duty-cycle bars from approximately 7 to 10 h