7/23/07

Joseph Harris

Hyuna Lee

Noo Li Jeon

UC Irvine

Dept of Biomedical Engineering

Laboratory of Dr. Noo Li Jeon

Fabrication of Neuron Device Protocol

Part 1:  Preparing the microfluidic neuron device

1.1)  A 3” silicon wafer with a pattern made out of SU-8 by photolithography is used to cast the Polydimethylsiloxane (PDMS) microfluidic mold.  We refer to the SU-8 patterned silicon wafer as master molds or “masters”.

1.2)  PDMS is mixed with curing agent at a w/w ratio of 10:1, that is, for 10 grams of PDMS, 1 gram of curing agent is added to it.

1.3)  The PDMS is then mixed well and poured onto the silicon wafer.  The silicon wafer is contained within a 10 cm polystyrene Petri dish.  It takes approximately 12 g to 15 g of PDMS to cover the master with a thickness of approximately 4 mm.

1.4)  The master with the PDMS is then placed in a vacuum dessicator with the lid off the Petri dish and vacuum is applied.  This is done to help remove air bubbles from the PDMS that were introduced during the stirring in of the curing agent.  It takes approximately 15 minutes for the air bubbles to be removed.

1.5)  The master with PDMS is removed from the dessicator.  An air gun with a 0.45 um filter is used to remove any remaining bubbles.

1.6)  The master is then placed on an 80˚C hot plate for 1 hour to cure.

If a vacuum dessicator is not available, the master can be left at room temperature for several hours.  The air bubbles will eventually dissipate on their own.  If no hot plate is available, the PDMS will cure at room temperature after 24 hours.

It is also important to make sure that the master is level during the curing process.

Part 2:  Cutting out the neuron microfluidic device

2.1)  Once the PDMS is cured on the master, it is taken to a clean hood.

2.2)  Using a surgical blade, a circle is cut around the perimeter of the devices.  It is important that when one inserts the blade into the PDMS to make contact with the silicon wafer but NOT put too much pressure on the wafer or else the master will crack.

2.3)  With a circle cut all the way around the devices (there are 4 devices per each 3” Silicon Master), the PDMS is carefully lifted off the master.  This can be started by gently twisting the surgical blade as it is inserted into the prior cut.

2.4)  Next, reservoirs are punched into the device using an 8 mm tissue biopsy punch.

2.5)  The devices are then quartered and excess PDMS trimmed away so that the device will fit neatly on Corning Cover Glass (No. 1) 24 mm X40 mm.  Nitrogen air is used to blow away any excess debris.  Stubborn debris can also be removed using 3M Scotch Brand 471 vinyl tape.

Part 3:  Plasma bonding the devices to glass cover slips.

3.1)  A Harrick Plasma Cleaner is used to bond the neuron devices to the cover slips.  Briefly, the devices are placed in the plasma cleaner and a vacuum pump used to evacuate air from the chamber.

3.2)  After the vacuum pressure has reached 300 milliTorr, power is turned on to the electrical coil and set to high.

3.3)  The electrical coil creates plasma, “a partially ionized gas consisting of electrons, ions and neutral atoms or molecules” - http://www.harrickplasma.com/plasma_physics.php, out of the remaining air molecules left in the chamber.  The plasma creates reactive species on the surfaces of the glass and the device which when placed together will form a permanent bond.  The plasma also turns the surfaces hydrophilic which helps facilitate the addition of liquids.  The plasma cleaner also sterilizes the devices.

The surface of the PDMS device that contains the microfluidic device is placed face up into the plasma cleaner along with the glass slides.

The plasma treated surfaces of the glass and microfluidic device are assembled in contact with each other in a clean bench, then placed in sterile 60 mm dishes.  The treat surfaces form a tight irreversible bond.

3.4)  Within 10 minutes of plasma bonding the devices to glass, Poly-L-Lysine (PLL) is added to the device.  After 10 minutes, the hydrophilicity of the device will decrease making it difficult for the PLL to enter the device.

Tech note:  It is important to check for air bubbles in the device after the addition of PLL.  Air bubbles in the main channel are undesirable as this would oxidize cells.  An easy method to remove any air bubbles present is to use a P200 pipet filled with 150 ul of liquid (in this case PLL), place the tip directly at the opening of the main channel, and depress the P200 forcefully.  This may need to be repeated several times but eventually the force from the pipet will drive the bubbles out.

3.5)  The devices containing PLL are allowed to incubate overnight in a standard tissue culture incubator at 37˚C with 5% CO2.

3.6)  The next day the devices are rinsed twice quickly with autoclaved dH2O.  During this process, liquid is never entirely removed from the main channels, only from the reservoirs.  Removing liquid from the main channels will introduce bubbles.  After (2) quick rinses, the devices are filled with autoclaved dH2O and placed back into the incubator for a minimum of 3 hours or over night.

3.7)  The next day, the devices are once again rinsed (2) times quickly with autoclaved dH2O.  Then neural basal media containing the necessary supplements B27 and glutamax for culturing primary rat neurons, is added to the devices.  The devices are placed back in the incubator for future use.  The devices are now ready for the seeding of neurons.

