Microfluidic Chips Controlled with Elastomeric Microvalve Arrays 
Nianzhen Li, Chris Sip, and Albert Folch
Department of Bioengineering, University of Washington, Seattle, WA, 98195, USA 

Context: 

Miniaturized microfluidic systems provide simple and effective solutions for low-cost point-of-care diagnostics and high-throughput biomedical assays.  Robust flow control and precise fluidic volumes are two critical requirements for these applications.  We have developed microfluidic chips featuring elastomeric polydimethylsiloxane (PDMS) microvalve arrays that: 1) need no extra energy source to close the fluidic path, hence the loaded device is highly portable; and 2) allow for microfabricating deep (up to 1 mm) channels with vertical sidewalls and resulting in very precise features.  

Here we demonstrate two such microfluidic chips for two different applications. The first chip allows for storing and mixing precise sub-nanoliter volumes of aqueous solutions at various mixing ratios.  The second chip allows for computer-controlled perfusion of microfluidic cell cultures.

The devices are easy to fabricate and simple to control.  Due to the biocompatibility of PDMS, these microchips could have broad applications in miniaturized diagnostic assays as well as basic biological studies. 
Protocol:

1. Microfluidic device design using CorelDraw or AutoCAD software.

2. Fabrication of silicon Masters using standard SU-8 photolithography.

3. Replica molding of PDMS from the master.  
4. Thin PDMS Membrane Manufacturing.

5. Multilayer PDMS device bonding and assembly. 
6. Computer-controlled opening and closing of PDMS microvalves by vacuum or pressure.
7. Parallel mixing of fluorescein with water in different defined nanoliter volumes. 

8. An integrated microfluidic system for computer-controlled perfusion of microfluidic cell cultures.

Detailed Protocol

1. Microfluidic device design using CorelDraw or AutoCAD software.

Principle of PDMS microvalves-based devices: The devices consist of three layers: a fluidic layer containing microchambers of various sizes, a “control layer” containing the microchannels necessary to actuate the fluidic path with microvalves, and a middle thin PDMS membrane that is bound to the control layer.  (show a picture of design side-view.) 

At rest, due to the compliance and hydrophobicity of PDMS, the membrane seals (reversibly) against its seat, therefore the chambers remain isolated from each other without energy input.  Valves can be opened by applying negative pressure (e.g. house vacuum), so the PDMS membrane deflects down and separates from the surface that supports the wall between two fluidic chambers, thus connecting the fluidic path.  Valve closure can be achieved by switching the pressure setting from vacuum to atmospheric pressure. 

Main advantages of our microvalve design:  1) No extra energy source is required to close the fluidic path, hence the loaded device is highly portable; and 2) the device can be built by PDMS replicas from photolithographically-patterned SU-8 molds, allowing for microfabricating deep (up to 1 mm) channels with vertical sidewalls (i.e. the height of the features can be specified independently of their width) and resulting in very precise features.
Next we will show the design of a parallel mixer that allows for storing and mixing precise sub-nanoliter volumes of aqueous solutions at various mixing ratios: 

The fluidic layer contains two arrays of microchambers: Along array A, the size of the microchambers decreases, starting from the left, from 200 m ( 400 m to 200 m ( 40 m; A10 is a 500 m ( 40 m chamber and is used just for fluidic connection in Array A; to the right of chamber A10 is a set of chambers symmetrically increasing in sizes.  The chambers in array B are designed such that the added volume of any two adjacent chambers in different rows always equal to. A0, A0r and B10, B10r are designed as respective controls for solutions A and B without mixing.

The control layer has two independently-controlled sets of valves.  A set of valves {V1} is used to connect two chamber arrays with their respective inlets, whereas a second set of valves {V2} is used to connect each pair of chambers in the two arrays. (show design in Corel file?)

Masks containing these designs were printed at high resolution (8000, to 20,000 dpi) on transparency films through commercial services.  (show masks?)

2. Fabrication of silicon Masters using standard SU-8 photolithography.

Standard SU-8 photolithography methods were used to create masters for the microfluidic layer and the valve control layer.  

To facilitate release, prior to PDMS replication the SU-8 master was silanized by exposure to a vapor of a fluorosilane ((tridecafluoro-1,1,2,2,-tetrahydrooctyl)-1-trichlorosilane) in house vacuum at room temperature for 30 min. 

 (We won’t show the photolithography procedures, since they are very time consuming and had to be done in a clean room.  But do we want to show the masters with SU-8 features?)  

3. Replica molding of PDMS from the master.  
Fluidic layer and control layer are made by replica molding of PDMS from SU-8 masters. 

Mixing thoroughly PDMS pre-polymer and cross-linker at 10:1 wt. ratio, de-bubble in a desiccator for 10-15 min until bubbles clear. 
Cut silicone tubings into 1-2 cm long pieces. Choose appropriate size of tubings according to applications. We use 1.14 mm I.D. tubings here for easy connection to 1/16 inch O.D tubings later.  

Use Duco-Cement to glue tubings onto the inlet regions of the SU-8 master of the control layer.  Be careful not to use too much glue.  Since the silicone tubings are made of the same components as PDMS, these tubings will be embedded into the PDMS microfluidic device, creating air and fluidic tight inlets/outlets.

In our device, inlet regions are designed on the masks of both the fluidic and the control layers, but silicone tubing inlets are molded only into one layer (for example, control layer) of the device.  To create inlets to the fluidic layer, we manually remove or puncture the few sections of the membrane that cover the inlet regions.  Therefore, after alignment and assembly, all microchannels (those that carry flow as well as those that control the valves) are accessible from the top of the device so that the bottom surface is planar, enabling imaging of the device on a conventional microscope stage.

Carefully pour de-bubbled PDMS onto both masters, around the tubings in the control-layer master.  De-bubble again in a desiccator.  After de-bubble is complete, put into 65 oC oven for > 1 hour for curing. 

Remove cured PDMS-covered masters from the oven.  

Cut individual devices from the masters (each master contains three identical devices) and peel off.

Remove glue from the inlet regions using a needle or a pair of forceps.

Take the control layer PDMS into the cleanroom. 

4. Thin PDMS Membrane Manufacturing.

As shown in the device principle, the middle layer consists of a ~12 μm-thick PDMS membrane.  

Mix 10:1 wt. ratio of PDMS prepolymer/curing agent mixture with hexane (3:1 wt. ratio) by vortexing. 

Move into a clean room. (A dust-free environment is critical for ensuring that the PDMS membranes are free of defects; dust particles may result in membranes containing holes and/or defectively bound to the replica mold.) 

Put a silanized 3 inch –diameter wafer onto the vacuum chuck of a Soletec spinner.  The wafer has to be silanized (derivaticed with fluorosilane) prior to PDMS spinning to facilitate the release of PDMS from silicon surfaces.  The Teflon bowl outside of the chuck was wrapped with plastic film for easy cleaning. 

Dispense 2-3 ml of PDMS/Hexane mixture onto the wafer using an 18-Gauge syringe needle (for minimizing bubbles). 

Set spin parameters. Spin at 7000 rpm for 30 sec, resulting in a PDMS film of ~ 12 μm thickness. 

Heat the wafer at 85 oC for 4 min on a hot plate to cure the PDMS film.  

5. Multilayer PDMS device bonding and assembly. 
Permanently bond the control layer and the PDMS membrane by putting both into O2 plasma chamber for 30 sec and then bring into contact. 

Wait for 5 min, and peel off the control layer with the membrane. 

Remove membranes on the inlets areas so that both layers are accessible from the top via tubings.  (In our device, inlet regions are designed on the masks of both the fluidic and the control layers, but silicone tubing inlets are molded only into one layer (for example, control layer) of the device.  To create inlets to the fluidic layer, we manually remove or puncture the few sections of the membrane that cover the inlet regions.  Therefore, after alignment and assembly, all microchannels (those that carry flow as well as those that control the valves) are accessible from the top of the device so that the bottom surface is planar, enabling imaging of the device on a conventional microscope stage.) 

Go back to main lab. Align the control layer (with tubings as inlets) with the fluidic layer (planar) under the stereoscope. (The scope is connected to a CCD camera, so it’s visible how alignment is achieved.)  Because PDMS seals onto PDMS, no permanent bonding is required. 

6. Computer-controlled opening and closing of PDMS microvalves by vacuum or pressure.

After device alignment and assembly, 1/16 inch O.D. (1/32 inch I.D.) Tygon tubings are inserted into the silicone inlets and connected to the pressure sources or the fluidic reservoirs.  

For opening and closing valves, pressures are controlled by a vacuum line and an air pressure line connected through two pressure regulators to an array of miniature three-way solenoid valves. 

The solenoid valves are connected to National Instruments data acquisition hardware controlled via Labview software (National Instruments). 

Device operation and membrane deflection are visualized with a CCD camera.  
7. Parallel mixing of fluorescein with water in different defined nanoliter volumes. 

To fill the microchambers, valve set {V1} is opened to allow the flow of two solutions to arrays A and B, respectively.  Flow of solutions can be achieved either by hand or by vacuum pulling controlled with solenoid valves.  If air bubbles form in the microchambers, more solution can be pushed to remove the bubbles, or the device can be left for a few minutes and bubbles will disappear due to the air permeability of PDMS. 

The valve set {V1} is then closed, thereby isolating each chamber in both arrays.  

Next, valve set {V2} is opened to allow fluid mixing between adjacent chambers in different arrays.  Mixing only takes ~ 1-2 min to complete for these volumes.  

Closing {V2} then pushes the fluid above back to each fluidic chamber and chambers deform back to their original shape.  Since the two fluidic arrays are designed with chambers of 11 different sizes, 11 different dilution ratios are produced in a single mixing step.  

The device is easy to fabricate and simple to control.  The volumes are photolithographically defined and, thus, very precise.  Fluid and reagents can be stored in the microdevice for several days, allowing for highly portable assays.  Notably, PDMS is biocompatible, so the device has broad applicability in miniaturized diagnostic assays as well as in cell-based assays such as drug screening and enzyme-based biomolecule detection.

8. An integrated microfluidic system for computer-controlled perfusion of microfluidic cell cultures.

The principal of our integrated microfluidic system: The device is capable of the automated perfusion of multiple chemicals solutions to a single cell culture chamber.  The inlets are controlled by microvalves which can be activated in any sequence of single inlets, various combinations, or all at once.  The device is capable of producing gradients or mixtures of the various solutions.    

This device also consists of three layers: a fluidic layer, a “control layer”, and a middle thin PDMS membrane.  The main differences are the method for connecting the inlets and an additional bonding step.   The inlets are connected by using 18 gauge needles which are inserted in to the PDMS through the control layer.  This allows for a denser packing of inlets than the silicone tubing.  The additional bonding step irreversibly binds the membrane to the fluid layer in addition to the control layer.

The features of our integrated microfluidic system:  The device is capable of automated perfusion of 16 different solutions to a cell culture chamber using a multiplexed valving scheme.  The channel design ensures that the resistance of all the inlets is balanced.  Our microvalve design isolates solutions and controls rinsing through integrated channels for the rapid removal of fluid which limits cross-contamination.  An integrated herringbone mixer can be activated to produce mixtures of different inlets.  Additionally, there are four varying resistance channels that can be activated to alter the flow rate.  
Alternative fabrication steps: After removing the PDMS control layer from the replica mold, holes are bored through the layer using a Harris micropunch with a 1.2 mm diameter.  This later allows for the easy insertion and removal of 18 gauge needles to connect vacuum lines and fluid reservoirs to the device.  The compliance of the PDMS provides a tight seal around the needles to effectively deliver fluid or vacuum pressure.  
The additional bonding step ensures a proper seal for the fluidic channels at higher pressures.  To bond the membrane/control layer to the fluid layer they are both placed into O2 plasma chamber for 45 sec at a lower power setting of 55.  The layers are carefully aligned under the stereoscope. Immediately after alignment, vacuum lines are inserted into the device to actuate the valves.  The step is critical because the valves will irreversibly bond closed due to the configuration of our valve scheme.   The valves are then turned on and off for at least a half hour to ensure that they do not bond closed or open.  

Operation of the device: We will now demonstrate the pressurized filling and operation of the device using colored dyes to simulate various solutions.
